## Room Names (hidden)
06/04/2022 DRAFT
AusHFG Room Code | AHFG Room name | HIU Recommended Room Name | Prog Area
1BR-BA | 1 Bed Room - Bariatric | Bed - Bariatric | 18
1BR-H-12 | 1 Bed Room - Holding, 12m2 | Bed - Holding | 12
1BR-IC | 1 Bed Room - Intensive Care | Bed - Intensive | 25
1BR-IS-N1 | 1 Bed Room - Isolation - Negative Pressure, Type 1 | Bed - Isolation Neg | 15
1BR-IS-N2 | 1 Bed Room - Isolation - Negative Pressure, Type 2 | Bed - Isolation Neg | 16.5
1BR-MH-A | 1 Bed Room - Mental Health - Inboard Ensuite, 15m2 | Bed - MH | 15
1BR-MH-C | 1 Bed Room - Mental Health - Back to Back Ensuites, 15m2 | Bed - MH | 15
1BR-SP-A1 | 1 Bed Room - Special, Inboard Ensuite, Type 1 | Bed - Special | 18
1BR-SP-A2 | 1 Bed Room - Special, Inboard Ensuite, Type 2 | Bed - Special | 18
1BR-SP-B | 1 Bed Room - CCU, 20m2 | Bed - CCU | 20
1BR-ST-A1 | 1 Bed Room - Inboard Ensuite, Type 1 | Bed 1P | 16
1BR-ST-A2 | 1 Bed Room - Inboard Ensuite, Type 2 | Bed 1P | 16.5
1BR-ST-A3 | 1 Bed Room - Inboard Ensuite, Type 3 | Bed 1P | 16.5
1BR-ST-B1 | 1 Bed Room - Outboard Ensuite, Type 1 | Bed 1P | 16.5
1BR-ST-B2 | 1 Bed Room - Outboard Ensuite, Type 2 | Bed 1P | 17.5
1BR-ST-B3 | 1 Bed Room - Outboard Ensuite, Type 3 | Bed 1P | 17.5
1BR-ST-D | 1 Bed Room - Back to Back Ensuite | Bed 1P | 17.5
2BR-ST-A1 | 2 Bed Room - Inboard Ensuite, Type 1 | Bed 2P | 27
2BR-ST-A2 | 2 Bed Room - Inboard Ensuite, Type 2 | Bed 2P | 29
2BR-ST-B | 2 Bed Room - Outboard Ensuite | Bed 2P | 28
4BR-ST | 4 Bed Room - Inboard Ensuite | Bed 4P | 58
ADLB | ADL Bathroom | ADL Bath | 12
ADLD | ADL Dining | ADL Dining | 12
ADLK | ADL Kitchen | ADL Kitchen | 12
ADLL | ADL Laundry | ADL Laundry | 8
AHBBF | After Hours Blood Fridge | A/H Blood Fridge | 3
AHDR | After Hours Drug Store | A/H Drug Store | 5
AIRLE-12 | Airlock - Entry, 12m2 | Airlock | 12
AIRLE-6 | Airlock - Entry, 6m2 | Airlock | 6
ANAE-16 | Anaesthetic Preparation Room, 16m² | Anaesthetic Prep | 16
ANRM | Anteroom | Anteroom | 6
AUD-CR | Audiology Control Room | Audiology Control | 9
AUD-TR | Audiology Testing Room | Audiology Testing | 14
BATH | Bathroom | Bathroom | 15
BATM-2 | Bay - ATM, 2m2 | Bay - ATM | 2
BBEV-ENC | Bay - Beverage, Enclosed, 5m2 | Bay - Bev | 5
BBEV-OP | Bay - Beverage, Open Plan, 4m2 | Bay - Bev | 4
BBW | Bay - Blanket/ Fluid Warmer | Bay - Blanket/ Fluid | 1
BES | Bay - Emergency Shower | Bay - Emergency Shower | 1
BFLW-ENC | Bay - Flowers, Enclosed, 4m2 | Bay - Flowers | 4
BFLW-OP | Bay - Flowers, Open Plan, 2m2 | Bay - Flowers | 2
BHW | Bay - Height/ Weight | Bay – Ht/Wt | 2
BHWS-A | Bay - Handwashing, Type A | Bay – HW A | 1
BHWS-B | Bay - Handwashing, Type B | Bay – HW B | 1
BIRM-A | Birthing Room - LDR Without Bath | Birthing - LDR | 34
BIRM-B | Birthing Room - LDR With Bath | Birthing - LDR (bath) | 40.5
BLIN | Bay - Linen | Bay - Linen | 2
BLST | Blood Store | Store - Blood | 2
BMEQ | Bay - Mobile Equipment | Bay - Mob Eq | 4
BMT-4 | Bay - Meal Trolley, 4m2 | Bay - Meal Trolley | 4
BPATH | Bay - Pathology | Bay - Path | 1
BPH | Bay - Public Telephone | Bay - Telephone | 2
BPROP | Bay - Property, Staff | Bay - Property Staff | 2
BPTS | Bay - Pneumatic Tube | Bay - PTS | 1
BRES | Bay - Resuscitation Trolley | Bay - Resus Trolley | 1.5
BVM-3 | Bay - Vending Machine | Bay - Vending | 3
BWC | Bay - Wheelchair Park | Bay - Wheelchair | 4
BWD-1 | Bay - Water Dispenser | Bay – Water | 1
CHPT | Change Cubicle - Patient, 2m2 | Change - Patient | 2
CHPT-12 | Change Room - Patient (Male/Female), 12m2 | Change - Patient (M/F) | 12
CHPT-D | Change Cubicle - Accessible, 4m2 | Change - Accessible | 4
CHST-10 | Change - Staff, 10m2 | Change - Staff | 10
CHST-35 | Change - Staff, 35m2 | Change - Staff | 35
CLAB-EP | Catheter Laboratory - Electrophysiology Studies | Cath Lab - Electrophysiology Studies | 60
CLAB-I | Catheter Laboratory - Interventional | Cath Lab - Interventional | 55.5
CLCR-EP | Catheter Laboratory EP Control Room | Cath Lab -EP Control Room | 17
CLCR-I | Catheter Laboratory Interventional Control Room | Cath Lab Interventional Control Room | 14
CLN-10 | Clean Store, 10m2 | Store - Cleaner | 10
CLN-MED-20 | Clean Store / Medication Room, 20m2 | Clean Store / Medication | 20
CLN-MED-S | Clean Store / Medication Room, Sub | Clean Store / Med - Sub | 8
CLRM-10 | Cleaner's Room, 10m2 | Cleaner | 10
CLRM-5 | Cleaner's Room, 5m2 | Cleaner | 5
CLUP-10 | Clean-Up Room - Shared, 10m2 | Clean-Up - Shared | 10
CLUP-7 | Clean-Up Room, 7m2 | Clean-Up | 7
CLUP-P | Clean-Up Room (Pathology), 12m2 | Clean-Up (Pathology) | 12
COMM | Communications Room | Comms | 12
CONS | Consult Room | Consult | 12
CONS-ENT-OP | Consult Room - ENT/ Ophthalmology | Consult - ENT/ Ophthalmology | 16
CONS-UN | Consult Room - Universal Access | Consult - Universal Access | 17
CORR | Corridor - Patient Treatment Areas | Corridor - Clinical | 0
CTCR | CT Imaging Control Room | CT Control | 12
CTIR | CT Imaging Room | CT Imaging | 45
CTPR | CT Planning Room | CT Planning | 45
DEN-MLB | Dental Laboratory, Minor | Dental Lab - Minor | 12
DENSR-1 | Dental Surgery, Type 1 | Dental Surgery | 14.5
DENSR-2 | Dental Surgery, Type 2 | Dental Surgery | 14.5
DINBEV-25 | Dining Room/Beverage Bay (Mental Health), 25m2 | Dining/Bev Bay - MH | 25
DINR | Dining - Patients | Dining - Patients | 72
DISP-10 | Disposal Room, 10m2 | Disposal | 10
DISP-8 | Disposal Room, 8m2 | Disposal | 8
DTUR-10 | Dirty Utility, 10m2 | Dirty Utility | 10
DTUR-12 | Dirty Utility, 12m2 | Dirty Utility | 12
DTUR-14 | Dirty Utility, 14m2 | Dirty Utility | 14
DTUR-S | Dirty Utility, Sub | Dirty Utility - Sub | 8
ECHO-TOE | Echocardiography - Transoesophageal | Echocardiography - Transoesophageal | 26
ECL-10 | Equipment Clean-Up, 10m2 | Clean-Up - Equip | 10
ECL-12 | Equipment Clean-Up/Loan Equipment, 12m2 | Clean-Up - Loan Equip | 12
ECL-14 | Equipment Clean-Up, 14m2 | Clean-Up - Equip | 14
ECL-8 | Equipment Clean-Up, 8m2 | Clean-Up - Equip | 8
ENPR | Procedure Room - Endoscopy | Procedure - Endoscopy | 45
ENS-ACC | Ensuite - Accessible, 7m2 | Ensuite - Access | 7
ENS-BA | Ensuite - Bariatric, 7m2 | Ensuite - Bariatric | 7
ENS-BR | Ensuite - Birthing, 7m2 | Ensuite - Birthing | 7
ENS-MH-A | Ensuite - Mental Health, Inboard, 5m2 | Ensuite - MH | 5
ENS-MH-B | Ensuite - Mental Health, Inboard Access from Corridor, 5m2 | Ensuite - MH | 5
ENS-SH | Ensuite - Shared, 6m2 | Ensuite - Shared | 6
ENS-SP | Ensuite - Special, 6m2 | Ensuite - Special | 6
ENS-ST-A1 | Ensuite - Inboard - Alternative 1, 5m2 | Ensuite | 5
ENS-ST-A2 | Ensuite - Inboard - Alternative 2, 5m2 | Ensuite | 5
ENS-ST-A3 | Ensuite - Inboard - Alternative 3, 5m2 | Ensuite | 5
ENS-ST-B | Ensuite - Outboard, 5m2 | Ensuite | 5
ENS-ST-C | Ensuite - Back to Back, 5m2 | Ensuite | 5
FLUO | Fluoroscopy Room | Fluoroscopy | 40
FLUOC | Fluoroscopy Control Room | Fluoroscopy Control | 12
GENXR | General X-Ray Room | General X-Ray | 38
GYAH-GP | Gymnasium, Group Therapy | Gymnasium - Group | 105
GYAH-ID | Gymnasium, Individual Treatment | Gymnasium - Individual | 76
HYDP | Hydrotherapy Pool | Pool - Hydrotherapy | 260
INTF-MH | Interview Room - Mental Health | Interview - MH | 14
INTV | Interview Room | Interview | 12
LAUN-MH | Laundry - Mental Health, 6m2 | Laundry - MH | 6
LAUN-PT | Laundry - Patient, 6m2 | Laundry - Patient | 6
LINAC | Linear Accelerator Treatment Room | Linear Accelerator - Treatment | 165
LINAC-CR | Linear Accelerator Control Room | Linear Accelerator - Control | 21
LNPA-12 | Lounge - Parent, 12m2 | Lounge - Parent | 12
LNPF-20 | Lounge - Patient / Family, 20m2 | Lounge - Patient / Family | 20
LNPT-10 | Lounge - Patient / Family, 10m2 | Lounge - Patient / Family | 10
LNPT-30 | Lounge - Patient / Family, 30m2 | Lounge - Patient / Family | 30
MAMMO | Mammography Room | Mammography | 16
MED-14 | Medication Room, 14m2 | Medication | 14
MEET-12 | Meeting Room, 12m2 | Meeting | 12
MEET-9 | Meeting Room, 9m2 | Meeting | 9
MEET-L-15 | Meeting Room, 15m2 | Meeting | 15
MEET-L-20 | Meeting Room, 20m2 | Meeting | 20
MEET-L-30 | Meeting Room, 30m2 | Meeting | 30
MEET-L-55 | Meeting Room, 55m2 | Meeting | 55
MOR-AU | Mortuary - Autopsy Room | Mortuary - Autopsy | 30
MOR-BH | Mortuary - Body Holding | Mortuary - Body Hold | 22
MOR-BR | Mortuary - Body Reception | Mortuary - Body Reception | 7
MOR-VR | Mortuary - Viewing Room | Mortuary - Viewing | 8
MOR-W | Mortuary - Waiting | Mortuary - Waiting | 9
MRICR | MRI Control Room | MRI Control | 14
MRIR | MRI Room | MRI | 46
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | Neonatal Bay - High Dep | 16.5
NBLD | Neonatal Bay - Low Dependency Care | Neonatal Bay - Low Dep | 12.5
OFF-2P | Office - 2 Person Shared, 12m2 | Office 2P | 12
OFF-3P | Office - 3 Person Shared, 15m2 | Office 3P | 15
OFF-4P | Office - 4 Person Shared, 20m2 | Office 4P | 20
OFF-CLN | Office - Clinical Workroom | Office - Clinical | 15
OFF-S12 | Office - Single Person, 12m2 | Office 1P | 12
OFF-S9 | Office - Single Person, 9m2 | Office 1P | 9
OFF-SWS | Office - Shared Workstation, 2.2m2 | Workstation | 2.2
OFF-WI-5 | Office - Write-up, 5m2 | Write-up | 5
OPG | OPG Room | OPG | 12
ORGN | Operating Room - General | Operating - General | 60
OVBR | Overnight Stay - Bedroom | O/S Bed | 10
OVES | Overnight Stay - Ensuite | O/S Ensuite | 5
PAR | Parenting Room | Parenting | 6
PBAT | Patient Bay, Emergency - Ambulance Triage | Patient Bay - Triage | 6.5
PBIC | Patient Bay - Intensive Care | Patient Bay - Intensive | 24.5
PBTR-A | Patient Bay, Emergency - Acute Treatment | Patient Bay - Acute | 12
PBTR-AS | Patient Room, Emergency - Acute Treatment Special | Patient ED - Acute Treatment Special | 15
PBTR-FT | Patient Bay, Emergency - Fast Track | Patient Bay ED - Fast Track | 6.5
PBTR-H-6 | Patient Bay - Holding, 6m2 | Patient Bay - Holding | 6
PBTR-H-9 | Patient Bay - Holding, 9m2 | Patient Bay - Holding | 9
PBTR-MD | Patient Bay - Medical Day Treatment | Patient Bay - Day | 9
PBTR-NA | Patient Bay, Emergency - Non Acute Treatment | Patient Bay ED - Non Acute | 10
PBTR-R | Patient Bay, Emergency - Resuscitation | Patient Bay ED - Resuscitation | 25
PBTR-RD-A | Patient Bay - Renal Dialysis, Chair | Patient Bay - Renal Dialysis, Chair | 9
PBTR-RD-B | Patient Bay - Renal Dialysis, Bed | Patient Bay - Renal Dialysis, Bed | 9
PBTR-RS1 | Patient Bay - Recovery, Stage 1, 9m2 | Patient Bay - Recovery, Stage 1, 9m2 | 9
PET-CT | PET-CT Imaging Room | PET-CT Imaging Room | 50
PHA-CO | Pharmacy - Counter | Pharmacy - Counter | 9
PHA-DB | Pharmacy - Distribution Workstation | Pharmacy - Distribution Workstation | 6
PHA-DS | Pharmacy - Dispensing Workstation | Pharmacy - Dispensing Workstation | 4
PHA-PR | Pharmacy - Preparation Room, Non-Aseptic | Pharmacy - Preparation Room, Non-Aseptic | 12
PHA-RE | Pharmacy - Returns Workstation | Pharmacy - Returns Workstation | 6
PLAP-10 | Play Area - Paediatric, 10m2 | Play Area - Paediatric, 10m2 | 10
PLAP-20 | Play Area - Paediatric, 20m2 | Play Area - Paediatric, 20m2 | 20
PLNT-WT | Plant - Water Treatment | Plant - Water Treatment | 18
PLST | Plaster Room | Plaster Room | 18
PROC | Procedure Room | Procedure Room | 20
PTRY | Pantry | Pantry | 8
REC-E | Reception, Emergency | Reception, Emergency | 20
RECL-10 | Reception/ Clerical, 10m2 | Reception/ Clerical, 10m2 | 10
RECL-12 | Reception/ Clerical, 12m2 | Reception/ Clerical, 12m2 | 12
RECL-15 | Reception/ Clerical, 15m2 | Reception/ Clerical, 15m2 | 15
REPR | Reporting Room | Reporting Room | 12
REPW | Reporting Workstation | Reporting Workstation | 6
SCRB-4 | Scrub Up, 4m2 | Scrub Up, 4m2 | 4
SECL | Seclusion Room | Seclusion Room | 14
SECR-10 | Security Room, 10m2 | Security Room, 10m2 | 10
SHAC | Shower - Accessible | Shower - Accessible | 4
SHPT | Shower - Patient | Shower - Patient | 4
SHST | Shower - Staff | Shower - Staff | 3
SPECC | Specimen Collection Bay | Specimen Collection Bay | 9
SPECT-CT | SPECT-CT Imaging Room | SPECT-CT Imaging Room | 48
SPECT-CTCR | SPECT-CT Control Room | SPECT-CT Control Room | 14
SPREC | Specimen Reception/ Sort/ Preparation | Specimen Reception/ Sort/ Preparation | 20
SRM-15 | Staff Room, 15m2 | Staff Room, 15m2 | 15
SRM-35 | Staff Room, 35m2 | Staff Room, 35m2 | 35
SSTN-10 | Staff Station, 10m2 | Staff Station, 10m2 | 10
SSTN-14 | Staff Station, 14m2 | Staff Station, 14m2 | 14
SSTN-20 | Staff Station, 20m2 | Staff Station, 20m2 | 20
STAD | Store - Accountable Drugs | Store - Accountable Drugs | 7.5
STBK-20 | Store - Bulk, 20m2 | Store - Bulk, 20m2 | 20
STBK-40 | Store - Bulk, 40m2 | Store - Bulk, 40m2 | 40
STCL | Store - Cleaners | Store - Cleaners | 12
STEQ-14 | Store - Equipment, 14m2 | Store - Equipment, 14m2 | 14
STEQ-20 | Store - Equipment, 20m2 | Store - Equipment, 20m2 | 20
STFS-10 | Store - Files, 10m2 | Store - Files, 10m2 | 10
STFS-20 | Store - Files, 20m2 | Store - Files, 20m2 | 20
STGN | Store - General | Store - General | 9
STPP | Store - Patient Property | Store - Patient Property | 8
STPS-10 | Store - Photocopy/ Stationery, 10m2 | Store - Photocopy/ Stationery, 10m2 | 10
STPS-8 | Store - Photocopy/ Stationery, 8m2 | Store - Photocopy/ Stationery, 8m2 | 8
STRT | Stress Testing | Stress Testing | 20
STSS-20 | Store - Sterile Stock, 20m2 | Store - Sterile Stock, 20m2 | 20
STSS-CC | Store - Sterile Stock, Central Core | Store - Sterile Stock, Central Core | 34
TRIAGE-1 | Triage Assessment Room, Emergency - Type 1 | Triage Assessment Room, Emergency - Type 1 | 13
TRIAGE-2 | Triage Assessment Room, Emergency - Type 2 | Triage Assessment Room, Emergency - Type 2 | 13
TRMT | Treatment Room | Treatment Room | 14
TRMT-HTS | Treatment Room - Hand Therapy / Splinting | Treatment Room - Hand Therapy / Splinting | 28
TRMT-POD | Treatment Room - Podiatry | Treatment Room - Podiatry | 14
ULTR | Ultrasound Room | Ultrasound Room | 14
ULTR-PR | Ultrasound Room - Procedures | Ultrasound Room - Procedures | 20
WAIT-10 | Waiting, 10m2 | Waiting, 10m2 | 10
WAIT-20 | Waiting, 20m2 | Waiting, 20m2 | 20
WAIT-30 | Waiting, 30m2 | Waiting, 30m2 | 30
WAIT-50 | Waiting, 50m2 | Waiting, 50m2 | 50
WAIT-SEC | Waiting - Secure, 6m2 | Waiting - Secure, 6m2 | 6
WAIT-SUB | Waiting - Sub, 5m2 | Waiting - Sub, 5m2 | 5
WCAC | Toilet - Accessible | Toilet - Accessible | 6
WCPT | Toilet - Patient | Toilet - Patient | 4
WCPU | Toilet - Public | Toilet - Public | 3
WCST | Toilet - Staff | Toilet - Staff | 3

## SoA Example Guide (hidden)
SoA Example Guide
2023-06-26
Objectives
- Provide the steps needed to develop a SoA from briefing stage, through design and into project delivery
- To provide for greater transparency across the development of project SoA reporting with particular focus on tracking changes against source AHFG HPU SoAs
- To achieve consistent reporting formats and outputs across the various projects
- To offer a template that can be adopted, mimicked or replicated
Assumptions
- The AHFG HPU SOA are used to establish and inform the briefing SoA
- The AHFG Standard Components will be used to populate project rooms and derived rooms
- Appropriate software will be used to manage the project data throughout the project design phases
- The project brief SoA will be digitally linked to the designed response
- All designed areas reported in the SoA will be calculated using the methodology described in AusHFG Part C
Establishing a Project SOA
- It is recommended that the AHFG SOA’s are used as a starting point
- If no direct match exists, apply the closest AHFG HPU match and keep record of the project actions to adjust the HPU in the SOA comments
- A record of the HPU revision, issue details and status should be kept at the point of project establishment
- The AHFG HPU comments should be captured and tagged with the SOA commencement date or ‘AHFG’
- The AHFG HPU SOA will need to have the number fields converted from text to number format
Circulation, Travel, Engineering and Facade
- Circulation is indicated for each sub-area in the AHFG HPU’s and these figures should be incorporated into the project SoA
- Travel and engineering briefed % figures should be agreed and incorporated on an SOA summary sheet and reported alongside the project totals
- A % can also be added to the summary page to represent the façade depending on project scale and type
AHFG HPUs and Standard components
(from AHFG: Part B - Health Facility Briefing and Planning 0090 - Standard Components)
Standard Components are a key feature of the AHFGs and provide detailed information on commonly used rooms and spaces across healthcare projects. Each Standard Component has an associated Room Data Sheet (RDS) and Room Layout Sheet (RLS). The Standard Components can be divided into two broad types: Components used across most Health Planning Units (HPU). Examples include a Toilet – Staff and Meeting Rooms. HPU specific components which are routinely only used to inform specialty requirements. Examples include an Operating Room and Dental Surgery. Standard Components can be used for a variety of purposes including mandatory compliance or as a starting point for design, however they have been developed with a view to providing:
• a shared understanding of room names, fit-out requirements and functions; and
• a starting point for design.
In practice, Standard Components provide a broad overview of requirements. In many circumstances, requirements will be influenced by model of care/services, technology and other areas of innovation. The recent review of standard components has reduced the number of room type variations i.e. waiting rooms at many sizes. Instead, only one or two examples have been retained. These rooms will provide a well-developed ‘starting point’ to develop alternate room sizes. For example, a waiting room or staff change room which will range in size depending on the service size and staffing numbers. These rooms will be known as Standard Component – Derived Rooms.
Room Names
- See the list of AHFG standard components with recommended room names attached
- Use key word for all rooms names (Toilet - Staff)
- Avoid the use of commas and apostrophes in room names as these will interfere with any csv generated reports
- It’s advisable to remove the word ‘room’ from all rooms to reduce unnecessary content
- Remove the area from the room names ahead of concept design
- Inpatient bedrooms can have the name reduced to ‘Bed’ and have the correct AHFG room code assigned to identify type
- Some software systems allow a room name to be connected with a room name tag. However for those that don’t, it will be necessary to reduce the room name length so that it will fit inside the planned room once the room link is established (example: Toilet -Patient can be reduced to WCPT in order for the label to fit inside the room when printed at 1:100)
Room types
Each HPU in the AHFGs contain a Schedule of Accommodation (SoA) that lists the room/space and indicates if a Standard Component or Standard Component – Derived Room is available. The names used in the SoA are consistent with Standard Component names so they are easily identifiable. In many specialised HPUs, Standard Components will not exist for all rooms/spaces listed in the SoA. Where a Standard Component is not available, these rooms will be broadly described in each HPU to provide guidance on:
- description and function;
- location and relationships; and
- other considerations.
The current Standard Components can be found at: www.healthfacilityguidelines.com.au/standardcomponents Part B - Health Facility Briefing and Planning Page 5 0090 - Standard Components, Revision 6.0, 01 March 2
The Project SoA should nominate each room type
- Standard component (SC)
- Standard Component Derived (SC-D)
- Non-standard component (NS)
Standard components
Each HPU in the AHFGs contain a Schedule of Accommodation (SoA) that lists the room/space and indicates if a Standard Component or Standard Component – Derived Room is available. The names used in the SoA are consistent with Standard Component names so they are easily identifiable. In many specialised HPUs, Standard Components will not exist for all rooms/spaces listed in the SoA. Where a Standard Component is not available, these rooms will be broadly described in each HPU to provide guidance on:
- description and function;
- location and relationships; and
- other considerations.
Where possible it is recommended that a standard component is used as a starting point for non-standard rooms. Refer to the list attached for non-standard room base data recommendations.
Comments section:
- It is helpful to include comments that track the movement of an SOA in a consistent format
- This means the information is widely accessible and can be included in the SOA report
- Using a consistent format that starts with the date means the data can be sorted by date
- A ‘when what who’ system is recommended (example: 211019 room moved from support TK)
- The date should be kept in a consistent format once the format is decided regular checks for consistency in the early phases of briefing is important
- The date format should exclude symbols and allow for logical sorting (example: 210203 for the 3rd of February 20201)
- Keeping the notes clear and concise is important however the overuse of acronyms is not recommended
- Always include the initials of who has made the change so that this data is regularly available
- Minimise the use of acronyms in the annotation but establish some commonly acceptable examples (I.E. UG1 = User Group number 1)
- IF a room has moved from one area to another - always comment where the room has moved from – as where it has moved to will be clearly apparent
- If a room has reduced in briefed area – always include the figure it has reduced from – as what it has been changes to will always be evident
- Keep the most recent comment at the top of the comments list by pushing older notes below – in this way the data is organised according to the timeline with the most recent being filterable in and xlsx reports
Corridors
- Circulation is a briefed percentage at the project SoA start-up
- As corridors are added to the design they should be added to the SoA
- The designed areas for these corridors should be compared with the briefed percentage values
- avoid creating a single space for a complex corridor layouts
- Corridors should be logically separated in the planning to ensure a room number appears in logical locations
Functional Areas
There are 3 functional area groups that should be captured across all HPU SOA reports
1. Procedure / Imaging Rooms – to include a count of each procedure room and a total for each department (Define procedure rooms to include Minor procedures, Cath Lab, Endoscopy, Operating Theatres and Imaging rooms). Exclude interview rooms and consult/exam rooms (and treatment rooms?)
2. Patient Spaces – to include a count for all patient occupiable bed or chair spaces for all departments – definition to include all overnight inpatient beds, AAU bed spaces, short stay bed spaces, recliner chair spaces, and including all recovery spaces in theatre suites, cath labs and endoscopy. For 4 bed rooms for example the count should be 4 patient spaces. Exclude toilets & en-suites (a list of what’s not included? Operating rooms? Mental Health de-escalation rooms, seclusion rooms)
3. Staff workstation spaces – to include a count for each staff occupiable workstation space in all offices,, workrooms, open plan office spaces. Definition of a workstation – any space where a staff member will sit or stand and may include a PC workstation (including or excluding operating theatre workstations? Control room spaces? hot desks? Storeroom workstations?)
Room templates
- The room templates can be set up or imported to represent the information contained within the AHFG standard components
- The template codes should match the AHFG Standard Component codes
- These templates should be used to track any variations proposed for the project
- Tracked variations to the templates will include changes to recommended room size, room description, finishes and FF&E
- Any bespoke project templates should be coded such that they are recognised easily (I.E. prefixed or suffixed with a number or letter)
Area variations
- SoA is to highlight any designed areas under or over briefed areas by 5%
- Red denotes areas under 5%
- Amber denotes area over 5%
Rooms added and deleted
- Additional rooms should have reasoning added to comments column
- Deleted rooms should be reported in SoAs with reasoning added to comments column
- Deleted rooms can be removed from SoA following each project milestone and SoA report
Rooms moved between sub-departments
- Include reasoning in Room Notes column
- deleted original room and add note to state where it has moved to
- add room to alternative department, adding a comment to indicate where it has moved from

## SoA Guide Cover Sheet
SoA Example Cover Sheet
2023-06-26
Objectives
- Provide the steps needed to develop a SoA from briefing stage, through design and into project delivery
- To provide for greater transparency across the development of project SoA reporting with particular focus on tracking changes against source AHFG HPU SoAs
- To achieve consistent reporting formats and outputs across the various projects
- To offer a template that can be adopted, mimicked or replicated
Assumptions
- The AHFG HPU SOA are used to establish and inform the briefing SoA
- The AHFG Standard Components will be used to populate project rooms and derived rooms
- Appropriate software will be used to manage the project data throughout the project design phases
- The project brief SoA will be digitally linked to the designed response
- All designed areas reported in the SoA will be calculated using the methodology described in AusHFG Part C

## 1-HPU VERSION REGISTER
Project X
AusHFG record of HPU status as implemented in the Project X
Project Department - AHFG HPU
Issue Date:
06/04/2022 DRAFT
PROJECT HPU CODE | PROJECT HPU NAME | AHFG HPU NAME | AHFG HPU Revision # | Date published | Status | Comments
AMU | Adult Acute Mental Health Inpatient Unit | HPU 360 Intensive Care Unit | 7 | 2019-07-01 | Current | <add any project specific comments on HPU adoption>
<add dept code> | <add project department name> | <add AHFG HPU name> | <add revision #> | <add publish date> |  | <add any project specific comments on HPU adoption>
<add dept code> | <add project department name> | <add AHFG HPU name> | <add revision #> | <add publish date> |  | <add any project specific comments on HPU adoption>
<add dept code> | <add project department name> | <add AHFG HPU name> | <add revision #> | <add publish date> |  | <add any project specific comments on HPU adoption>

## DO NOT EDIT (hidden)
Project X - DO NOT EDIT - HIU TO UPDATE
Building & Department Summary of Briefed vs. Designed
06/04/2022 DRAFT
|  |  |  | Building A | Building B | Building C | Building D
|  | Total Gross Dept Area (GDA) |  | 1269.11
|  | Travel |  | 228.4398
|  | Engineering |  | 317.2775
|  | Gross Building Area(GBA) |  | 1814.8273 | 0 | 0 | 0
Façade
|  | Total GBA plus Façade |  | 1814.8273 | 0 | 0 | 0
Percentage Factors
|  | Travel |  | 18%
|  | Engineering |  | 25%
|  | Façade |  | 0%
Functional Unit Count
Building | Department Code | Department | Briefed Area | Area Incl. Briefed Circ. | Designed Area | Area Incl. Designed Circ. | Diff. m2 | Diff. % (+/-5%) | Briefed External | Designed External | Room Count | Procedure Rooms | Patient Spaces | Staff W/Spaces
EXAMPLE SHEET | ICU | Intensive Care Department | 1042.9 | 1409.975 | 1132.11 | 1269.11 | -140.865 | -9.991% | 50 | 45 | 92 | 0 | 15 | 13
Your project | AMU | Adult Acute Mental Health
Building A Total |  |  | 1042.9 | 1409.975 | 1132.11 | 1269.11 | -140.865 | -9.991% | 50 | 45 | 92 | 0 | 15 | 13
<Building Name> | <add dept code> | <add project department name>
Building B Total |  |  | 0 | 0 | 0 | 0 |  |  | 0 | 0 | 0 | 0 | 0 | 0
Grand Total |  |  | 1042.9 | 1409.975 | 1132.11 | 1269.11 | -140.865 | -9.991% | 50 | 45 | 92 | 0 | 15 | 13

## 2-HPU SOA for RECORD
FOR PROJECT RECORD AND FUTURE REFERENCE ONLY
Download a copy of each relevant HPU (one per xls sheet) - refer to SOA Technical Guidance Note
Australasian Health Facility Guidelines
Part B - Health Facility Briefing and Planning
HPU 0360 Intensive Care Unit, Revision 7.0, July 2019 (FOR EXAMPLE)
ENTRY, WAITING & FAMILY SUPPORT AREAS
32.00 Support Areas | AusHFG Room Code | Room / Space | SC / SC-D | PICU - 14 beds |  | Adult ICU - 14 beds |  | Adult ICU - 56 beds (4 pods) |  | Remarks | Total GDA (20 Bed) | #REF!
32.00 Support Areas |  |  |  | Qty | m2 | Qty | m2 | Qty | m2 |  | Approx. GDA/Bed | #REF!
32.00 Support Areas | RECL-12 | Reception / Clerical | Yes | 1 | 10 (o) | 1 | 10 (o) | 1 | 12 | Optional. Provision will depend on the size of the unit and staffing arrangements.
32.00 Support Areas | WAIT-10 | Waiting | Yes | 1 | 15 | 1 | 15 | 1 | 40 | Nominal area. To be calculated at 1.2m2 per able- bodied person, 1.5m2 per wheelchair occupant.
32.00 Support Areas | PLAP-10 | Play Area - Paediatric | Yes | 1 | 10 (o) | Shared |  | 1 | 10 (o) | Optional.
32.00 Support Areas | BVM-3 | Bay - Vending Machine | Yes | 1 | 2 | 1 | 2 | 1 | 4 | Locate close to but not within waiting area given level of noise and heat generation.
32.00 Support Areas | BWD-1 | Bay - Water Dispenser | Yes | 1 | 1 | 1 | 1 | 1 | 1
32.00 Support Areas | BBEV-OP | Bay - Beverage, Open Plan | Yes |  |  |  |  | 1 | 4
32.00 Support Areas | WCPU-3 | Toilet - Public | Yes | 1 | 3 | 1 | 3 | 4 | 3
32.00 Support Areas | WCAC | Toilet- Accessible | Yes | 1 | 6 (o) | 1 | 6 (o) | 1 | 6 (o) | Optional. Include if not located nearby.
| PAR | Parenting Room | Yes | 1 | 6 (o) | 1 | 6 (o) | 1 | 6 (o) | Optional. Include if not located nearby.
| INTF | Interview Room | Yes | 1 | 14 | 1 | 14 | 3 | 14 | For interviews with relatives. Include dual egress. Consider including some interview rooms within the clinical zone.
| MEET-L-20 | Meeting Room | Yes | 1 | 20 (o) | 1 | 20 (o) | 1 | 20 | Multifunctional meeting room for family and staff meetings. Larger meeting room frequently required for end of life discussions.
| OVBR | Overnight Stay - Bedroom | Yes | 2 | 10 | 1 | 10 | 3 | 10 | For families. Provision depends on local operational policies. May be designed for use during the day as overflow interview rooms
| OVES | Overnight Stay - Ensuite | Yes | 2 | 4 | 1 | 4 | 3 | 4
| PROP-2 | Property Bay | Yes | 1 | 1 (o) | 1 | 1 (o) | 1 | 2 (o) | Optional. For family / visitor belongings if using overnight stay rooms as multipurpose spaces.
|  | Discounted Circulation % |  |  | 25% |  | 25% |  | 25%
PATIENT AREAS
| AusHFG Room Code | Room / Space | SC / SC-D | PICU - 14 beds |  | Adult ICU - 14 beds |  | Adult ICU - 56 beds (4 pods) |  | Remarks
|  |  |  | Qty | m2 | Qty | m2 | Qty | m2
| 1BR-IC | 1 Bed Room - Intensive Care | Yes | 7 | 25 | 13 | 25 | 50 | 25 | The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space.
| PBIC | Patient Bay - Intensive Care | Yes | 6 | 24.5 |  | 24.5 |  | 24.5
| 1BR-IC | 1 Bed Room - Intensive Care (Class N Isolation) | Yes | 1 | 25 | 1 | 25 | 4 | 25 | No. dependent on patient cohort and local jurisdictional policies.
| 1BR-IC | 1 Bed Room - Intensive Care (Class P Isolation) | Yes |  | 25 (o) |  | 25 (o) | 2 | 25 (o) | Optional. Depends on cohort and local jurisdictional requirements.
| ANRM | Anteroom | Yes | 1 | 6 | 1 | 6 | 4 | 6 | Attached to 1 Bed Rooms - Isolation Class N (neg pressure ventilation).
| BATH | Bathroom | Yes | 1 | 15 | 1 | 15 | 2 | 15 | Requires gases. May be shared between 2 pods subject to design.
| ENS-SP | Ensuite - Special | Yes | 2 | 6 | 2 | 6 | 8 | 6 | Ensuites dedicated to negative pressure, positive pressure and standard isolation rooms. Number dependent on isolation room requirements. PICUs should include bathrooms with baths for younger children.
| ENS-SP | Ensuite - Special | Yes | 3 | 6 | 3 | 6 | 12 | 6 | Ensuites to be shared between remaining rooms (1:4 ratio). PICUs should include bathrooms with baths for younger children.
| SSTN-20 | Staff Station | Yes | 1 | 25 | 1 | 25 | 4 | 25 | 1 central staff station per pod.
| OFF-CLN | Office – Clinical Workroom | Yes | 1 | 15 | 1 | 15 | 4 | 15 | 1 per pod.
|  | Decentralised Staff Workstation |  | 4 | 2 | 7 | 2 | 28 | 2 | 1 between 2 enclosed rooms.
| BHWS-A | Bay - Handwashing, Type A | Yes | 2 | 1 | 2 | 1 | 8 | 1 | Also included within each patient bay.
| BLIN | Bay - Linen | Yes | 1 | 2 | 1 | 2 | 4 | 2 | 1 per pod.
| BRES | Bay - Resuscitation | Yes | 1 | 1.5 | 1 | 1.5 | 4 | 1.5 | 1 per pod.
| BPATH | Bay - Pathology | Yes | 1 | 3 | 1 | 3 | 4 | 3 | 1 per pod. May be shared between 2 pods depending on layout.
| BPTS | Bay - Pneumatic Tube | Yes | 1 | 1 | 1 | 1 | 4 | 1 | 1 per pod.
| BMEQ-4 | Bay - Mobile Equipment | Yes | 3 | 4 | 3 | 4 | 12 | 4 | 3 per pod. Locate in quiet low traffic areas.
| BBW | Bay - Blanket / Fluid Warming | Yes | 1 | 1 | 1 | 1 | 2 | 1 | Shared between 2 pods if collocated, subject to design.
|  | Discounted Circulation % |  |  | 40% |  | 40% |  | 40%
SUPPORT AREAS
| AusHFG Room Code | Room / Space | SC / SC-D | PICU - 14 beds |  | Adult ICU - 14 beds |  | Adult ICU - 56 beds (4 pods) |  | Remarks
|  |  |  | Qty | m2 | Qty | m2 | Qty | m2
| BBEV-OP | Bay - Beverage, Open Plan | Yes | 1 | 4 | 1 | 4 | 2 | 4 | Share between 2 pods where collocated.
| BMT-4 | Bay - Meal Trolley | Yes | 1 | 4 (o) | 1 | 4 (o) | 2 | 4 (o) | Optional. Depends on local operational policy for food services. Share between 2 pods where collocated.
|  | Formula Room |  | 1 | 9 |  |  |  |  | Used for formula preparation and storage, and sterilising of expressing equipment within PICUs.
| STDR-10 | Medication Room | Yes | 1 | 12 | 1 | 12 | 2 | 20 | Share between 2 pods where collocated. Area to be distributed if automated dispensing units are provided centrally within central staff area.
|  | Store - Clean |  | 1 | 30 | 1 | 30 | 2 | 60 | Clean stock. Distribution between pods dependent on layout.
| STEQ-20 | Store - Equipment | Yes | 1 | 28 | 1 | 28 | 1 | 112 | 2m2 per bed. May include specialty stores eg for retrieval equipment, renal, allied health etc.
| STGN-9 | Store - General | Yes | 1 | 14 | 1 | 14 | 2 | 28
|  | Respiratory / Biomedical Workroom |  | 1 | 12 | 1 | 12 | 1 | 20 | For set up and servicing.
| DTUR-12 | Dirty Utility | Yes | 1 | 12 | 1 | 12 | 4 | 12 | 1 per pod.
| CLUP-7 | Clean-Up Room | Yes | 1 | 7 | 1 | 7 | 1 | 24
| CLRM-5 | Cleaner's Room | Yes | 1 | 5 | 1 | 5 | 4 | 5 | Number and distribution depends on layout.
| DISP-10 | Disposal Room | Yes | 1 | 8 | 1 | 8 | 1 | 20
|  | Simulation Room |  | 1 | 25 (o) |  |  | 1 | 25 (o) | Optional. Requirement will depend on education and training requirements / access to central simulation facilities.
|  | Simulation Control Room | [5 empty] | 1 | 10 (o) | Optional.
|  | Workroom - Telehealth |  | 1 | 12 (o) | 1 | 12 (o) | 1 | 12 | Optional. Inclusion depends on operational guideline of unit. Can be used for viewing medical images and discussions, and for telehealth activities for liaison regarding transfers, virtual rounds etc.
| WCST | Toilet - Staff | Yes | 2 | 3 | 2 | 3 | 6 | 3 | For ease of access from clinical areas.
|  | Discounted Circulation % |  |  | 40% |  | 40% |  | 40%
STAFF AREAS
| AusHFG Room Code | Room / Space | SC / SC-D | PICU - 14 beds |  | Adult ICU - 14 beds |  | Adult ICU - 56 beds (4 pods) |  | Remarks
|  |  |  | Qty | m2 | Qty | m2 | Qty | m2
| OFF-S9 | Office - Single Person, 9m2 | Yes |  | 9 |  | 9 |  | 9 | Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas.
| OFF-S12 | Office - Single Person, 12m2 | Yes |  | 12 |  | 12 |  | 12 | Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas.
|  | Office - Workstation, 4.4m2 |  |  | 4.4 |  | 4.4 |  | 4.4 | Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas.
|  | Office - Workstation, 5.5m2 |  |  | 5.5 |  | 5.5 |  | 5.5 | Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas.
| MEET-L-15 | Meeting Room | Yes | 1 | 15 | 1 | 15 | 1 | 12
| MEET-L-20 | Meeting Room | Yes | Shared |  | Shared |  | 1 | 20 | Assume shared access to 20m2 meeting room under Entry / Waiting & Family Support for smaller services.
| MEET-L-55 | Meeting Room | Yes |  |  |  |  | 1 | 50
| STPS-8 | Store - Photocopy / Stationery | Yes | 1 | 5 | 1 | 5 | 2 | 8
| SRM-18 | Staff Room | Yes | 1 | 18 | 1 | 18 | 1 | 60
| CHST-10 | Change - Staff (Male / Female) | Yes | 1 | 10 | 1 | 10 | 1 | 40 | Includes toilets, showers, lockers; size depends on the staffing per shift. (Female)
| CHST-10 | Change - Staff (Male / Female) | Yes | 1 | 8 | 1 | 8 | 1 | 32 | Includes toilets, showers, lockers; size depends on staffing per shift. (Male)
| WCAC | Toilet - Accessible | Yes | Shared |  | Shared |  | 1 | 6
| OVBR | Overnight Stay - Bedroom | Yes | 1 | 10 (o) | 1 | 10 (o) | 1 | 10 (o) | Optional. Requirement depends on staffing arrangements. Close access to toilet / shower facilities required.
|  | Discounted Circulation % |  |  | 25% |  | 25% |  | 25%

## 3-SoA Summary
SUMMARY - Project X EXAMPLE
Schedule of Accommodation - Concept Design
Building A
Issue Date: 28/10/2021
Version 2.1
No | Department | Project Brief |  | Designed Area Concept Design |  |  | Functional Unit Count Concept Design |  |  | Remarks / Room Notes
|  | AHFG HPU SoA Area | Total Dept Brief Area | Total Dept Designed Area | Diff. m2 | Diff. % (+/-5%) | Procedure / Treatment Spaces | Patient Bed / Chair Spaces | Staff W/Station Spaces
1 | Intensive Care Unit | 1305.875 | 1409.975 | 1269.11 | -140.865 | -9.991% | 0 | 15 | 20 | Room sizes refined and developed.
Circ. |  | 346.975 | 367.075 | 320 | -47.075 | -12.82%
2 | Example Department A | 1200 | 1287 | 1312 | 25 | 1.943% | 2 | 24 | 16 | Example only (demonstrate SoA Summary)
Circ. |  | 410 | 412 | 450 | 38 | 9.223% |  |  |  | Cultural rooms added in response to NZ DGN Example only (demonstrate SoA Summary)
3 | Example Department B | 1780 | 1818 | 1890 | 72 | 3.96% | 8 | 36 | 18 | Example only (demonstrate SoA Summary)
Circ. |  | 520 | 535 | 540 | 5 | 0.9346%
4 | Example Department C | 2272 | 2309 | 2280 | -29 | -1.256% | 10 | 18 | 22 | Example only (demonstrate SoA Summary)
Circ. |  | 603 | 660 | 675 | 15 | 2.273%
Total Circ |  |  | 1974.075 | 1985 | -61.94 | 0.5534%
BUILDING AREA GRAND TOTAL |  |  | 6823.975 | 6751.11 | -72.865 | -1.068%
Total Plant 11% |  |  | 750.6372 | 742.6221 | -8.0152 | -1.068% |  |  |  | 11% of Grand Total - GFA (AHFG Part C - one level building)
Functional Unit Total | [6 empty] | 20 | 93 | 76
EXTERNAL AREA DEPT TOTALS | Intensive Care Unit | 0 | 50 | 65 | 15 | 30% |  |  |  | Staff outdoor and patient terrace areas added
| Department A | 20 | 20 | 20 |  | 0% |  |  |  | Example only (demonstrate SoA Summary)
| Department B | 20 | 20 | 20 |  | 0% |  |  |  | Example only (demonstrate SoA Summary)
| Department C | 20 | 20 | 20 |  | 0% |  |  |  | Example only (demonstrate SoA Summary)
| EXTERNAL AREA GRAND TOTAL | 60 | 110 | 125 | 15 | 13.64%

## 4-SoA Example - Brief Dept 1
Project X EXAMPLE - BRIEF SoA
Schedule of Accommodation - Briefing Stage
Building A
Intensive Unit
(Reference:HPU 0360 Intensive Care Unit, Revision 7.0, July 2019)
Issue Date: | 2021-10-28
|  |  |  | Project Brief | [7 empty] | Functional Unit Count
Sub Area | Room Name | AHFG Code / Project Code | SC / SC-D / NS | Room Count | AHFG Area | Brief Area | Total Area | Circ % | Designed Area | Diff. m2 | Diff. % (+/-5%) | Procedure / Treatment Spaces | Patient Bed / Chair Spaces | Staff W/Station Spaces | Remarks / Room Notes
ENTRY, WAITING & FAMILY SUPPORT AREAS | Reception / Clerical |  |  |  | 10 | 0 | 0 |  | 13.9 | 13.9 |  |  |  |  | 211029 UG1 deleted as service model doesn’t require JB 210801 AHFG Optional. Provision will depend on the size of the unit and staffing arrangements
| Waiting | WAIT-10 | SC | 1 | 15 | 20 | 20 |  | 34.7 | 14.7 | 73.5% |  |  |  | 211029 reduced from 40m2 UG1 JB 210801 AHFG Nominal area. To be calculated at 1.2m2 per able- bodied person, 1.5m2 per wheelchair occupant
| Play Area - Paed | PLAP-10 | SC | 1 | 10 | 10 | 10 |  | 9 | -1 | -10% |  |  |  | 210801 AHFG Optional
| Bay - Vending | BVM-3 | SC | 1 | 2 | 2 | 2 |  | 2.7 | 0.7 | 35% |  |  |  | 211029 UG1 further reduced from 3m2 SG 210903 single vending machine so reduced from 4m2 AG 210801 AHFG Locate close to but not within waiting area given level of noise and heat generation
| Bay - Water Disp | BWD-1 | SC | 1 | 1 | 1 | 1 |  | 0.95 | -0.05 | -5%
| Bay - Bev | BBEV-OP | SC | 1 | 4 | 5 | 5 |  | 10.75 | 5.75 | 115% |  |  |  | 210925 increased from 4m2 for recycling station JB
| Toilet - Public | WCPU-3 | SC | 1 | 3 | 3 | 3 |  | 3.5 | 0.5 | 16.67%
| Toilet- Access | WCAC | SC | 1 | 6 | 6 | 6 |  | 5 | -1 | -16.67% |  |  |  | 211029 reduced from 2 as another located nearby JB 210801 AHFG Optional. Include if not located nearby
| Parenting | PAR | SC | 1 | 6 | 6 | 6 |  | 4 | -2 | -33.33% |  |  |  | 210801 AHFG Optional. Include if not located nearby
| Interview | INTF | SC | 1 | 14 | 12 | 12 |  | 12.3 | 0.3 | 2.5% |  |  |  | 211029 UG1 reduced from 14m2 JB 210801 AHFG For interviews with relatives. Include dual egress. Consider including some interview rooms within the clinical zone
| Whanau | PRJ-WH | NS | 1 | 0 | 18 | 18 |  | 18 |  | 0% |  |  |  | 211030 UG2 room added AG
| Meeting | MEET-L-20 | SC-D | 1 | 20 | 18 | 18 |  | 18.5 | 0.5 | 2.778% |  |  |  | 211102 UG3 reduced from 20m2 12-14 max occupants JB 210801 AHFG Multifunctional meeting room for family and staff meetings. Larger meeting room frequently required for end of life discussions
| Bed - Overnight | OVBR | SC | 2 | 10 | 12 | 24 |  | 30 | 6 | 25% |  |  |  | 211029 UG1 increased from 10m2 for small study desk and bed store JB 210801 AHFG For families. Provision depends on local operational policies. May be designed for use during the day as overflow interview rooms
| Ensuite - Overnight | OVES | SC | 2 | 4 | 4 | 8 |  | 9.6 | 1.6 | 20%
| Bay - Prop | PROP-2 | SC | 1 | 2 | 2 | 2 |  | 2.75 | 0.75 | 37.5% |  |  |  | 210801 AHFG Optional. For family / visitor belongings if using overnight stay rooms as multipurpose spaces
Sub Total | [6 empty] | 135 |  | 175.65 | 40.65 | 30.11%
Circ. | [6 empty] | 33.75 | 25% | 43.9125
PATIENT AREAS | 1 Bed - Intensive Care | 1BR-IC | SC | 13 | 25 | 25 | 325 |  | 324 | -1 | -0.3077% |  | 13 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Pt Bay - Intensive Care | PBIC | SC |  | 24.5 |  | 0 |  | 0 | [5 empty] | 211029 UG1 number TBC in Concept Design JB
| 1 Bed Room - Intensive Care (Class N Isolation) | 1BR-IC | SC | 1 | 25 | 25 | 25 |  | 24.8 | -0.2 | -0.8% |  | 1 |  | 210801 AHFG No. dependent on patient cohort and local jurisdictional policies
| 1 Bed - Intensive (Class P Isolation) | 1BR-IC | SC | 1 | 25 | 25 | 25 |  | 24.8 | -0.2 | -0.8% |  | 1 |  | 210801 AHFG Optional. Depends on cohort and local jurisdictional requirements
| Ante | ANRM | SC | 1 | 6 | 6 | 6 |  | 4.3 | -1.7 | -28.33% |  |  |  | 210801 AHFG Attached to 1 Bed Rooms - Isolation Class N (neg pressure ventilation)
| Bath | BATH | SC | 1 | 15 | 16 | 16 |  | 15.1 | -0.9 | -5.625% |  |  |  | 210801 AHFG Requires gases. May be shared between 2 pods subject to design
| Ensuite - Special | ENS-SP | SC | 2 | 6 | 7 | 14 |  | 14.5 | 0.5 | 3.571% |  |  |  | 210801 AHFG Ensuites dedicated to negative pressure, positive pressure and standard isolation rooms. Number dependent on isolation room requirements. PICUs should include bathrooms with baths for younger children
| Ensuite - Special | ENS-SP | SC | 3 | 6 | 6 | 18 |  | 16.3 | -1.7 | -9.444% |  |  |  | 210801 AHFG Ensuites to be shared between remaining rooms (1:4 ratio). PICUs should include bathrooms with baths for younger children
| Staff Station | SSTN-20 | SC | 1 | 25 | 20 | 20 |  | 20.23 | 0.23 | 1.15% |  |  | 5 | 210801 AHFG 1 central staff station per pod
| Office – Clinical | OFF-CLN | SC | 1 | 15 | 17 | 17 |  | 14.95 | -2.05 | -12.06% |  |  | 4 | 210801 AHFG 1 per pod
| Office - WS |  | NS | 7 | 2 | 1.5 | 10.5 |  | 10.25 | -0.25 | -2.381% |  |  | 7 | 210801 AHFG 1 between 2 enclosed rooms
| Bay - HW - A | BHWS-A | SC | 2 | 1 | 1 | 2 |  | 8 | 6 | 300% |  |  |  | 210801 AHFG Also included within each patient bay
| Bay - Linen | BLIN | SC | 1 | 2 | 2 | 2 |  | 8.4 | 6.4 | 320% |  |  |  | 210801 AHFG 1 per pod
| Bay - Resus | BRES | SC | 1 | 1.5 | 1.5 | 1.5 |  | 6.5 | 5 | 333.3% |  |  |  | 210801 AHFG 1 per pod
| Bay - Path | BPATH | SC | 1 | 3 | 3 | 3 |  | 4.9 | 1.9 | 63.33% |  |  |  | 210801 AHFG 1 per pod. May be shared between 2 pods depending on layout
| Bay - PTS | BPTS | SC | 1 | 1 | 1 | 1 |  | 1.5 | 0.5 | 50% |  |  |  | 210801 AHFG 1 per pod
| Bay - Mob Equip | BMEQ-4 | SC | 3 | 4 | 5 | 15 |  | 12 | -3 | -20% |  |  |  | 210801 AHFG 3 per pod. Locate in quiet low traffic areas
| Bay - Blank / Fluid | BBW | SC | 1 | 1 | 1 | 1 |  | 1.2 | 0.2 | 20% |  |  |  | 210801 AHFG Shared between 2 pods if collocated, subject to design
| <add additional rooms here> | [13 empty] | *date-what/why/reasoning-initials*
| <add additional rooms here> | [13 empty] | *date-what/why/reasoning-initials*
| <add additional rooms here> | [13 empty] | *date-what/why/reasoning-initials*
Sub Total | [6 empty] | 502 |  | 511.73 | 9.73 | 1.938%
Circ. | [6 empty] | 200.8 | 40% | 204.692
SUPPORT AREAS | Bay - Bev | BBEV-OP | SC | 1 | 4 | 3 | 3 |  | 4.1 | 1.1 | 36.67% |  |  |  | 210801 AHFG Share between 2 pods where collocated
| Bay - Meal Trolley | BMT-4 | SC | 1 | 4 | 4 | 4 |  | 3.8 | -0.2 | -5% |  |  |  | 210801 AHFG Optional. Depends on local operational policy for food services. Share between 2 pods where collocated
| Formula |  |  | 1 | 9 | 0 | 0 |  | 0 | [5 empty] | 211029 UG21 deleted - not required JB 210801 AHFG Used for formula preparation and storage, and sterilising of expressing equipment within PICUs
| Medication | STDR-10 | SC | 1 | 12 | 12 | 12 |  | 15 | 3 | 25% |  |  |  | 210801 AHFG Share between 2 pods where collocated. Area to be distributed if automated dispensing units are provided centrally within central staff area
| Store - Clean |  | NS | 1 | 30 | 33 | 33 |  | 27 | -6 | -18.18% |  |  |  | 210801 AHFG Clean stock. Distribution between pods dependent on layout
| Store - Equip | STEQ-20 | SC | 1 | 28 | 30 | 30 |  | 35 | 5 | 16.67% |  |  |  | 211029 increased from 112m2 for 3rd access door JB 210801 AHFG 2m2 per bed. May include specialty stores e.g. for retrieval equipment, renal, allied health etc
| Store - General | STGN-9 | SC | 1 | 14 | 15 | 15 |  | 14.9 | -0.1 | -0.6667% |  |  |  | 211029 increased from 28m2 for additional shelving UG1 JB
| Respiratory / Biomedical |  | NS | 1 | 12 | 35 | 35 |  | 32.3 | -2.7 | -7.714% |  |  |  | 210801 AHFG For set up and servicing
| Dirty Utility | DTUR-12 | SC | 1 | 12 | 12 | 12 |  | 12.23 | 0.23 | 1.917% |  |  |  | 210801 AHFG 1 per pod
| Clean-Up | CLUP-7 | SC | 1 | 7 | 7 | 7 |  | 7.9 | 0.9 | 12.86%
| Cleaners | CLRM-5 | SC | 1 | 5 | 5 | 5 |  | 5.2 | 0.2 | 4% |  |  |  | 210801 AHFG Number and distribution depends on layout
| Disposal | DISP-10 | SC | 1 | 8 | 8 | 8 |  | 7.35 | -0.65 | -8.125%
| Simulation |  | NS | 1 | 25 | 25 | 25 |  | 0 | -25 | - |  |  |  | 220306 room required for educational purposes TK 210801 AHFG Optional. Requirement will depend on education and training requirements / access to central simulation facilities
| Simulation Control |  |  | 0 | 10 | 0 | 0 |  | 0 | [5 empty] | 220306 deleted as not required as less than 56 bed unit TK 210801 AHFG Optional
| Workroom - Telehealth |  | NS | 1 | 12 | 12 | 12 |  | 12.1 | 0.1 | 0.8333% |  |  |  | 210801 AHFG Optional. Inclusion depends on operational guideline of unit. Can be used for viewing medical images and discussions, and for telehealth activities for liaison regarding transfers, virtual rounds etc.
| Toilet - Staff | WCST | SC | 2 | 3 | 3 | 6 |  | 6.22 | 0.22 | 3.667% |  |  |  | 210801 AHFG For ease of access from clinical areas
| Comms | COMM | SC | 1 | 14 | 14 | 14 | [7 empty] | Communication Room
Sub Total | [6 empty] | 207 |  | 183.1 | -23.9 | -11.55%
Circ. | [6 empty] | 82.8 | 40% | 73.24
STAFF AREAS | Office - 1P | OFF-S9 | SC | 1 | 9 | 9 | 9 |  | 10.1 | 1.1 | 12.22% |  |  | 3 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Office - 1P | OFF-S12 | SC | 1 | 12 | 12 | 12 |  | 10.95 | -1.05 | -8.75% |  |  | 1 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Office - WS |  | NS | 1 | 4.4 | 4.4 | 4.4 |  | 4 | -0.4 | -9.091% |  |  | 3 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Office - WS |  | NS | 1 | 5.5 | 5.5 | 5.5 |  | 5.65 | 0.15 | 2.727% |  |  | 3 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Meeting | MEET-L-15 | SC | 1 | 15 | 15 | 15 |  | 16.1 | 1.1 | 7.333% |  |  |  | 211029 UG2 increased from 12m2 to include telehealth function JB
| Meeting | MEET-L-20 | SC | 1 | 20 | 20 | 20 |  | 18.8 | -1.2 | -6% |  |  |  | 210801 AHFG Assume shared access to 20m2 meeting room under Entry / Waiting & Family Support for smaller services
| Store - Photo / Stationery | STPS-8 | SC-D | 1 | 5 | 5 | 5 |  | 14.8 | 9.8 | 196% |  |  |  | 211029 Project wide decision - reduced from 8m2 room to open bay JB
| Staff Room | SRM-18 | SC | 1 | 18 | 60 | 60 |  | 50 | -10 | -16.67%
| Change - St (M / F) | CHST-10 | SC | 1 | 10 | 40 | 40 |  | 40 |  | 0% |  |  |  | 210801 AHFG Includes toilets, showers, lockers; size depends on the staffing per shift. (Female)
| Change - St (M / F) | CHST-10 | SC | 1 | 8 | 30 | 30 |  | 32 | 2 | 6.667% |  |  |  | 211029 reduced from 32m2 less lockers required JB 210801 AHFG Includes toilets, showers, lockers; size depends on staffing per shift. (Male)
| Toilet - Access | WCAC | SC | 1 | 6 | 6 | 6 |  | 6.3 | 0.3 | 5%
| Bed - Overnight | OVBR | SC | 1 | 10 | 10 | 10 |  | 12 | 2 | 20% |  |  |  | 210801 AHFG Optional. Requirement depends on staffing arrangements. Close access to toilet / shower facilities required
Sub Total | [6 empty] | 211.9 |  | 220.7 | 8.8 | 4.153%
Circ. | [6 empty] | 52.975 | 25% | 55.175
CORRIDORS | Corridor (placeholder) | CORR | SC-D | 1
Sub Total | [6 empty] | 370.325 |  | 0 |  | -
Total Area | [6 empty] | 1055.9 |  | 1091.18
Total Circulation | [6 empty] | 370.325 |  | 377.0195
Grand Total | [6 empty] | 1426.225 |  | 1468.1995 | 412.2995 | 2.943%
Total Plant
Functional Unit Total | [11 empty] | 0 | 15 | 26
EXTERNAL | Terrace - ICU Bed |  | NS | 1 |  | 50 |  |  | 45 | [5 empty] | 210801 Added. Ability for patients to have fresh air and light

## 5-SoA Example - Designed Dept 1
Project X EXAMPLE - DESIGNED SoA
Schedule of Accommodation - Concept Design
Building A
Intensive Care Unit
(Reference:HPU 0360 Intensive Care Unit, Revision 7.0, July 2019)
Issue Date: | 2021-10-28
[6 empty] | Project Brief |  |  |  | Concept Design |  |  | Functional Unit Count
Sub Area | Room Name | AHFG Code / Project Code | SC / SC-D / NS | Room count | Room Number | AHFG Area | Project Brief Area | Total Area | Circ % | Designed Area | Diff. m2 | Diff. % (+/-5%) | Procedure / Treatment Spaces | Patient Bed / Chair Spaces | Staff W/Station Spaces | Remarks / Room Notes
ENTRY, WAITING & FAMILY SUPPORT AREAS | Reception / Clerical |  |  | 1 |  | 10 | 0 | 0 |  | 0 | [5 empty] | 211029 UG1 deleted as service model doesn’t require JB 210801 AHFG Optional. Provision will depend on the size of the unit and staffing arrangements
| Waiting | WAIT-10 | SC | 1 | 10022 | 15 | 20 | 20 |  | 34.7 | 14.7 | 73.5% |  |  |  | 211029 reduced from 40m2 UG1 JB 210801 AHFG Nominal area. To be calculated at 1.2m2 per able- bodied person, 1.5m2 per wheelchair occupant
| Play Area - Paed | PLAP-10 | SC | 1 | 10023 | 10 | 10 | 10 |  | 9 | -1 | -10% |  |  |  | 210801 AHFG Optional
| Bay - Vending | BVM-3 | SC | 1 | 10024 | 2 | 2 | 2 |  | 2.7 | 0.7 | 35% |  |  |  | 211029 UG1 further reduced from 3m2 SG 210903 single vending machine so reduced from 4m2 AG 210801 AHFG Locate close to but not within waiting area given level of noise and heat generation
| Bay - Water Disp | BWD-1 | SC | 1 | 10025 | 1 | 1 | 1 |  | 0.95 | -0.05 | -5%
| Bay - Bev | BBEV-OP | SC | 1 | 10026 | 4 | 5 | 5 |  | 10.75 | 5.75 | 115% |  |  |  | 210925 increased from 4m2 for recycling station JB
| Toilet - Public | WCPU-3 | SC | 1 | 10027 | 3 | 3 | 3 |  | 3.5 | 0.5 | 16.67%
| Toilet- Access | WCAC | SC | 1 | 10028 | 6 | 6 | 6 |  | 5 | -1 | -16.67% |  |  |  | 211029 reduced from 2 as another located nearby JB 210801 AHFG Optional. Include if not located nearby
| Parenting | PAR | SC | 1 | 10029 | 6 | 6 | 6 |  | 4 | -2 | -33.33% |  |  |  | 210801 AHFG Optional. Include if not located nearby
| Interview | INTF | SC | 1 | 10030 | 14 | 12 | 12 |  | 12.3 | 0.3 | 2.5% |  |  |  | 211029 UG1 reduced from 14m2 JB 210801 AHFG For interviews with relatives. Include dual egress. Consider including some interview rooms within the clinical zone
| Meeting | MEET-L-20 | SC-D | 1 | 10033 | 20 | 18 | 18 |  | 18.5 | 0.5 | 2.778% |  |  |  | 211102 UG3 reduced from 20m2 12-14 max occupants JB 210801 AHFG Multifunctional meeting room for family and staff meetings. Larger meeting room frequently required for end of life discussions
| Bed - Overnight | OVBR | SC | 1 | 10034 | 10 | 12 | 12 |  | 12 |  | 0% |  |  |  | 211029 UG1 increased from 10m2 for small study desk and bed store JB 210801 AHFG For families. Provision depends on local operational policies. May be designed for use during the day as overflow interview rooms
| Bed - Overnight | OVBR | SC | 1 | 10035 | 10 | 12 | 12 |  | 12 |  | 0% |  |  |  | 211029 UG1 increased from 10m2 for small study desk and bed store JB 210801 AHFG For families. Provision depends on local operational policies. May be designed for use during the day as overflow interview rooms
| Ensuite - Overnight | OVES | SC | 1 | 10036 | 4 | 4 | 4 |  | 4.2 | 0.2 | 5%
| Ensuite - Overnight | OVES | SC | 1 | 10037 | 4 | 4 | 4 |  | 5 | 1 | 25%
| Bay - Prop | PROP-2 | SC | 1 | 10038 | 2 | 2 | 2 |  | 2.75 | 0.75 | 37.5% |  |  |  | 210801 AHFG Optional. For family / visitor belongings if using overnight stay rooms as multipurpose spaces
| Whanau | PRJ-WH | NS | 1 | 10031 | NA | 18 | 18 |  | 18 |  | 0% |  |  |  | 211030 UG2 room added AG
| Whanau | PRJ-WH | NS | 1 | 10032 | NA | 18 | 18 |  | 22 | 4 | 22.22% |  |  |  | 211030 UG2 room added AG
Sub Total | [7 empty] | 117 |  | 177.35 | 60.35 | 51.58%
Circ. | [7 empty] | 29.25 | 25% | 44.3375 |  | 51.58%
PATIENT AREAS | Bed - Intensive Care | 1BR-IC | SC | 1 | 20010 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20011 | 25 | 25 | 25 |  | 26 | 1 | 4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20012 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20013 | 25 | 25 | 25 |  | 24 | -1 | -4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20014 | 25 | 25 | 25 |  | 24 | -1 | -4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20015 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20016 | 25 | 25 | 25 |  | 28 | 3 | 12% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20017 | 25 | 25 | 25 |  | 24 | -1 | -4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20018 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20019 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20020 | 25 | 25 | 25 |  | 28 | 3 | 12% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20021 | 25 | 25 | 25 |  | 24 | -1 | -4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20022 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Pt Bay – Intensive Care | PBIC | SC |  | 20023 | 24.5 |  | 0 |  | 0 | [5 empty] | 211029 UG1 number TBC in Concept Design JB
| Bed - Intensive Care (N) | 1BR-IC | SC | 1 | 20024 | 25 | 25 | 25 |  | 24.8 | -0.2 | -0.8% |  | 1 |  | 210801 AHFG No. dependent on patient cohort and local jurisdictional policies
| Bed - Intensive Care (P) | 1BR-IC | SC | 1 | 20025 | 25 | 25 | 25 |  | 24.8 | -0.2 | -0.8% |  | 1 |  | 210801 AHFG Optional. Depends on cohort and local jurisdictional requirements
| Ante | ANRM | SC | 1 | 20026 | 6 | 6 | 6 |  | 6 |  | 0% |  |  |  | 210801 AHFG Attached to 1 Bed Rooms - Isolation Class N (neg pressure ventilation)
| Bath | BATH | SC | 1 | 20027 | 15 | 16 | 16 |  | 15.1 | -0.9 | -5.625% |  |  |  | 210801 AHFG Requires gases. May be shared between 2 pods subject to design
| Ensuite - Special | ENS-SP | SC | 1 | 20028 | 6 | 7 | 7 |  | 7 |  | 0% |  |  |  | 210801 AHFG Ensuites dedicated to negative pressure, positive pressure and standard isolation rooms. Number dependent on isolation room requirements. PICUs should include bathrooms with baths for younger children
| Ensuite - Special | ENS-SP | SC | 1 | 20029 | 6 | 7 | 7 |  | 6.9 | -0.1 | -1.429% |  |  |  | 210801 AHFG Ensuites dedicated to negative pressure, positive pressure and standard isolation rooms. Number dependent on isolation room requirements. PICUs should include bathrooms with baths for younger children
| Ensuite - Special | ENS-SP | SC | 1 | 20030 | 6 | 6 | 6 |  | 6 |  | 0% |  |  |  | 210801 AHFG Ensuites to be shared between remaining rooms (1:4 ratio). PICUs should include bathrooms with baths for younger children
| Ensuite - Special | ENS-SP | SC | 1 | 20031 | 6 | 6 | 6 |  | 5.9 | -0.1 | -1.667% |  |  |  | 210801 AHFG Ensuites to be shared between remaining rooms (1:4 ratio). PICUs should include bathrooms with baths for younger children
| Ensuite - Special | ENS-SP | SC | 1 | 20032 | 6 | 6 | 6 |  | 6.2 | 0.2 | 3.333% |  |  |  | 210801 AHFG Ensuites to be shared between remaining rooms (1:4 ratio). PICUs should include bathrooms with baths for younger children
| Staff Station | SSTN-20 | SC | 1 | 20033 | 25 | 20 | 20 |  | 20.23 | 0.23 | 1.15% |  |  | 5 | 210801 AHFG 1 central staff station per pod
| Office – Clinical | OFF-CLN | SC | 1 | 20034 | 15 | 17 | 17 |  | 14.95 | -2.05 | -12.06% |  |  | 4 | 210801 AHFG 1 per pod
| Office - WS |  | NS | 1 | 20035 | 2 | 1.5 | 1.5 |  | 1.3 | -0.2 | -13.33% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20036 | 2 | 1.5 | 1.5 |  | 1.5 |  | 0% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20037 | 2 | 1.5 | 1.5 |  | 1.8 | 0.3 | 20% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20038 | 2 | 1.5 | 1.5 |  | 1.8 | 0.3 | 20% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20039 | 2 | 1.5 | 1.5 |  | 1.3 | -0.2 | -13.33% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20040 | 2 | 1.5 | 1.5 |  | 1.5 |  | 0% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20041 | 2 | 1.5 | 1.5 |  | 1.5 |  | 0% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Bay - HW - A | BHWS-A | SC | 1 | 20042 | 1 | 1 | 1 |  | 1.2 | 0.2 | 20% |  |  |  | 210801 AHFG Also included within each patient bay
| Bay - HW - A | BHWS-A | SC | 1 | 20043 | 1 | 1 | 1 |  | 0.9 | -0.1 | -10% |  |  |  | 210801 AHFG Also included within each patient bay
| Bay - Linen | BLIN | SC | 1 | 20044 | 2 | 2 | 2 |  | 8.4 | 6.4 | 320% |  |  |  | 210801 AHFG 1 per pod
| Bay - Resus | BRES | SC | 1 | 20045 | 1.5 | 1.5 | 1.5 |  | 6.5 | 5 | 333.3% |  |  |  | 210801 AHFG 1 per pod
| Bay - Path | BPATH | SC | 1 | 20046 | 3 | 3 | 3 |  | 4.9 | 1.9 | 63.33% |  |  |  | 210801 AHFG 1 per pod. May be shared between 2 pods depending on layout
| Bay - PTS | BPTS | SC | 1 | 20047 | 1 | 1 | 1 |  | 1.5 | 0.5 | 50% |  |  |  | 210801 AHFG 1 per pod
| Bay - Equip | BMEQ-4 | SC | 1 | 20048 | 4 | 5 | 5 |  | 6 | 1 | 20% |  |  |  | 210801 AHFG 3 per pod. Locate in quiet low traffic areas
| Bay - Equip | BMEQ-4 | SC | 1 | 20049 | 4 | 5 | 5 |  | 4 | -1 | -20% |  |  |  | 210801 AHFG 3 per pod. Locate in quiet low traffic areas
| Bay - Equip | BMEQ-4 | SC | 1 | 20050 | 4 | 5 | 5 |  | 2 | -3 | -60% |  |  |  | 210801 AHFG 3 per pod. Locate in quiet low traffic areas
| Bay - Blank / Fluid | BBW | SC | 1 | 20051 | 1 | 1 | 1 |  | 1.2 | 0.2 | 20% |  |  |  | 210801 AHFG Shared between 2 pods if collocated, subject to design
Sub Total | [7 empty] | 502 |  | 513.18 | 11.18 | 2.227%
Circ. | [7 empty] | 200.8 | 40% | 205.272
SUPPORT AREAS | Bay - Bev | BBEV-OP | SC | 1 | 30026 | 4 | 3 | 3 |  | 4.1 | 1.1 | 36.67% |  |  |  | 210801 AHFG Share between 2 pods where collocated
| Bay - Meal Trolley | BMT-4 | SC | 1 | 30027 | 4 | 4 | 4 |  | 3.8 | -0.2 | -5% |  |  |  | 210801 AHFG Optional. Depends on local operational policy for food services. Share between 2 pods where collocated
| Formula |  |  | 0 |  | 9 | 0 | 0 |  | 0 | [5 empty] | 211029 UG21 deleted - not required JB 210801 AHFG Used for formula preparation and storage, and sterilising of expressing equipment within PICUs
| Medication | STDR-10 | SC | 1 | 30029 | 12 | 12 | 12 |  | 15 | 3 | 25% |  |  |  | 210801 AHFG Share between 2 pods where collocated. Area to be distributed if automated dispensing units are provided centrally within central staff area
| Store - Clean |  | NS | 1 | 30030 | 30 | 33 | 33 |  | 27 | -6 | -18.18% |  |  |  | 210801 AHFG Clean stock. Distribution between pods dependent on layout
| Store - Equip | STEQ-20 | SC | 1 | 30031 | 28 | 30 | 30 |  | 35 | 5 | 16.67% |  |  |  | 211029 increased from 112m2 for 3rd access door JB 210801 AHFG 2m2 per bed. May include specialty stores e.g. for retrieval equipment, renal, allied health etc
| Store - General | STGN-9 | SC | 1 | 30032 | 14 | 15 | 15 |  | 14.9 | -0.1 | -0.6667% |  |  |  | 211029 increased from 28m2 for additional shelving UG1 JB
| Respiratory / Biomedical |  | NS | 1 | 30033 | 12 | 35 | 35 |  | 32.3 | -2.7 | -7.714% |  |  |  | 210801 AHFG For set up and servicing
| Dirty Utility | DTUR-12 | SC | 1 | 30034 | 12 | 12 | 12 |  | 12.23 | 0.23 | 1.917% |  |  |  | 210801 AHFG 1 per pod
| Clean-Up | CLUP-7 | SC | 1 | 30035 | 7 | 7 | 7 |  | 7.9 | 0.9 | 12.86%
| Cleaner | CLRM-5 | SC | 1 | 30036 | 5 | 5 | 5 |  | 5.2 | 0.2 | 4% |  |  |  | 210801 AHFG Number and distribution depends on layout
| Disposal | DISP-10 | SC | 1 | 30037 | 8 | 8 | 8 |  | 7.35 | -0.65 | -8.125%
| Simulation |  | NS | 1 | 30038 | 25 | 25 | 25 |  | 23 | -2 | -8% |  |  |  | 220306 room required for educational purposes TK 210801 AHFG Optional. Requirement will depend on education and training requirements / access to central simulation facilities
| Control - Simulation |  |  | 1 |  | 0 | 0 | 0 |  | 0 | [5 empty] | 220306 deleted as not required as less than 56 bed unit TK 210801 AHFG Optional
| Workroom - Telehealth |  | NS | 1 | 30040 | 12 | 12 | 12 |  | 12.1 | 0.1 | 0.8333% |  |  |  | 210801 AHFG Optional. Inclusion depends on operational guideline of unit. Can be used for viewing medical images and discussions, and for telehealth activities for liaison regarding transfers, virtual rounds etc.
| Toilet - Staff | WCST | SC | 1 | 30041 | 3 | 3 | 3 |  | 3.1 | 0.1 | 3.333% |  |  |  | 210801 AHFG For ease of access from clinical areas
| Toilet - Staff | WCST | SC | 1 | 30042 | 3 | 3 | 3 |  | 2.9 | -0.1 | -3.333% |  |  |  | 210801 AHFG For ease of access from clinical areas
| Comms Room | COMM | SC | 1 | 30043 | 0 | 14 | 14 |  | 15 | 1 | 7.143%
| <add additional rooms here> | PRJ-XX
Sub Total | [7 empty] | 207 |  | 220.88 | 13.88 | 6.705%
Circ. | [7 empty] | 82.8 | 40% | 88.352
STAFF AREAS | Office - 1P | OFF-S9 | SC | 1 | 40001 | 9 | 9 | 9 |  | 10.1 | 1.1 | 12.22% |  |  | 1 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Office - 1P | OFF-S12 | SC | 1 | 40002 | 12 | 12 | 12 |  | 10.95 | -1.05 | -8.75% |  |  | 1 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Office - WS | OFF-WS | NS | 1 | 40003 | 4.4 | 4.4 | 4.4 |  | 4 | -0.4 | -9.091% |  |  | 1 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Office - WS | OFF-WS | NS | 1 | 40004 | 5.5 | 5.5 | 5.5 |  | 5.65 | 0.15 | 2.727% |  |  | 1 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Meeting | MEET-L-15 | SC | 1 | 40005 | 15 | 15 | 15 |  | 16.1 | 1.1 | 7.333% |  |  |  | 211029 UG2 increased from 12m2 to include telehealth function JB
| Meeting | MEET-L-20 | SC | 1 | 40006 | 20 | 20 | 20 |  | 18.8 | -1.2 | -6% |  |  |  | 210801 AHFG Assume shared access to 20m2 meeting room under Entry / Waiting & Family Support for smaller services
| Store - Photo /Stationery | STPS-8 | SC-D | 1 | 40007 | 5 | 5 | 5 |  | 14.8 | 9.8 | 196% |  |  |  | 211029 Project wide decision - reduced from 8m2 room to open bay JB
| Staff Room | SRM-18 | SC | 1 | 40008 | 18 | 60 | 60 |  | 50 | -10 | -16.67%
| Change - F | CHST-10 | SC | 1 | 40009 | 10 | 40 | 40 |  | 40 |  | 0% |  |  |  | 210801 AHFG Includes toilets, showers, lockers; size depends on the staffing per shift. (Female)
| Change - M | CHST-10 | SC | 1 | 40010 | 8 | 30 | 30 |  | 32 | 2 | 6.667% |  |  |  | 211029 reduced from 32m2 less lockers required JB 210801 AHFG Includes toilets, showers, lockers; size depends on staffing per shift. (Male)
| Toilet - Access | WCAC | SC | 1 | 40011 | 6 | 6 | 6 |  | 6.3 | 0.3 | 5%
| Bed - Overnight | OVBR | SC | 1 | 40012 | 10 | 10 | 10 |  | 12 | 2 | 20% |  |  |  | 210801 AHFG Optional. Requirement depends on staffing arrangements. Close access to toilet / shower facilities required
Sub Total | [7 empty] | 216.9 |  | 220.7 | 3.8 | 1.752%
Circ. | [7 empty] | 54.225 | 25% | 55.175
CORRIDORS | Corridor | PRJ-COR | SC-D | 1 | 60002 |  |  |  |  | 42 | [5 empty] | *add designed area only - leave brief m2 blank as this is briefed as a %*
| Corridor | PRJ-COR | SC-D | 1 | 60005 |  |  |  |  | 60 | [5 empty] | *add designed area only - leave brief m2 blank as this is briefed as a %*
| Corridor | PRJ-COR | SC-D | 1 | 60008 |  |  |  |  | 35 | [5 empty] | *add designed area only - leave brief m2 blank as this is briefed as a %*
PLANT | DB board | PRJ-PLT | NS | 1 | 70001 |  |  |  |  | 2 | [5 empty] | *i.e comms, FHR, risers, small localised plant room - leave brief m2 blank as this is briefed as a % (excluding Comms Room)*
| Service Riser 1 | PRJ-PLT | NS | 1 | 70002 |  |  |  |  | 2.5 | [5 empty] | *i.e comms, FHR, risers, small localised plant room - leave brief m2 blank as this is briefed as a % (excluding Comms Room)*
| FHR (fire hose reel) | PRJ-PLT | NS | 1 | 70003 |  |  |  |  | 0.5
| Plant Room | PRJ-PLT | NS | 1 | 70004 |  |  |  |  | 150
Total Circ | [6 empty] | Brief | 367.075 | Designed | 137 | 128.42 | -62.68%
Total Area - Rooms | [6 empty] | Brief | 1042.9 | Designed | 1132.11 | 256.84 | 8.554%
Grand Total | [7 empty] | 1409.975 |  | 1269.11 | 226.21 | -9.991%
Total Plant 11% | [6 empty] | Brief | 155.0973 | Designed | 155 | -0.09725 | -0.0627% |  |  |  | 11% of Grand Total - GFA
Functional Unit Total | [12 empty] | 0 | 15 | 20
EXTERNAL | Terrace - ICU Bed |  | NS | 1 | 90000 | 0 | 50 | 50 |  | 45 | -5 | -10% |  |  |  | 210801 Added. Ability for patients to have fresh air and light JC
| Staff Outdoor area |  | NS | 1 | 90001 | 0 | 0 | 0 |  | 20 | 20 |  |  |  |  | 210912 Added. Staff relaxation area JC
Grand Total | [7 empty] | 50 |  | 65 | 65 | 30%

## AHFG Calcs_Summary (hidden)
Project X EXAMPLE - DESIGNED SoA
Schedule of Accommodation - Concept Design
Building A
Intensive Care Unit
(Reference:HPU 0360 Intensive Care Unit, Revision 7.0, July 2019)
Issue Date: | 2021-10-28
[6 empty] | Project Brief |  |  |  | Concept Design |  |  | Functional Unit Count
Sub Area | Room Name | AHFG Code / Project Code | SC / SC-D / NS | Room count | Room Number | AHFG Area | Project Brief Area | Total Area | Circ % | Designed Area | Diff. m2 | Diff. % (+/-5%) | Procedure / Treatment Spaces | Patient Bed / Chair Spaces | Staff W/Station Spaces | Remarks / Room Notes
ENTRY, WAITING & FAMILY SUPPORT AREAS | Reception / Clerical |  |  | 1 |  | 10 | 0 | 0 |  | 0 | [5 empty] | 211029 UG1 deleted as service model doesn’t require JB 210801 AHFG Optional. Provision will depend on the size of the unit and staffing arrangements
| Waiting | WAIT-10 | SC | 1 | 10022 | 15 | 20 | 20 |  | 34.7 | 14.7 | 73.5% |  |  |  | 211029 reduced from 40m2 UG1 JB 210801 AHFG Nominal area. To be calculated at 1.2m2 per able- bodied person, 1.5m2 per wheelchair occupant
| Play Area - Paed | PLAP-10 | SC | 1 | 10023 | 10 | 10 | 10 |  | 9 | -1 | -10% |  |  |  | 210801 AHFG Optional
| Bay - Vending | BVM-3 | SC | 1 | 10024 | 2 | 2 | 2 |  | 2.7 | 0.7 | 35% |  |  |  | 211029 UG1 further reduced from 3m2 SG 210903 single vending machine so reduced from 4m2 AG 210801 AHFG Locate close to but not within waiting area given level of noise and heat generation
| Bay - Water Disp | BWD-1 | SC | 1 | 10025 | 1 | 1 | 1 |  | 0.95 | -0.05 | -5%
| Bay - Bev | BBEV-OP | SC | 1 | 10026 | 4 | 5 | 5 |  | 10.75 | 5.75 | 115% |  |  |  | 210925 increased from 4m2 for recycling station JB
| Toilet - Public | WCPU-3 | SC | 1 | 10027 | 3 | 3 | 3 |  | 3.5 | 0.5 | 16.67%
| Toilet- Access | WCAC | SC | 1 | 10028 | 6 | 6 | 6 |  | 5 | -1 | -16.67% |  |  |  | 211029 reduced from 2 as another located nearby JB 210801 AHFG Optional. Include if not located nearby
| Parenting | PAR | SC | 1 | 10029 | 6 | 6 | 6 |  | 4 | -2 | -33.33% |  |  |  | 210801 AHFG Optional. Include if not located nearby
| Interview | INTF | SC | 1 | 10030 | 14 | 12 | 12 |  | 12.3 | 0.3 | 2.5% |  |  |  | 211029 UG1 reduced from 14m2 JB 210801 AHFG For interviews with relatives. Include dual egress. Consider including some interview rooms within the clinical zone
| Meeting | MEET-L-20 | SC-D | 1 | 10033 | 20 | 18 | 18 |  | 18.5 | 0.5 | 2.778% |  |  |  | 211102 UG3 reduced from 20m2 12-14 max occupants JB 210801 AHFG Multifunctional meeting room for family and staff meetings. Larger meeting room frequently required for end of life discussions
| Bed - Overnight | OVBR | SC | 1 | 10034 | 10 | 12 | 12 |  | 12 |  | 0% |  |  |  | 211029 UG1 increased from 10m2 for small study desk and bed store JB 210801 AHFG For families. Provision depends on local operational policies. May be designed for use during the day as overflow interview rooms
| Bed - Overnight | OVBR | SC | 1 | 10035 | 10 | 12 | 12 |  | 12 |  | 0% |  |  |  | 211029 UG1 increased from 10m2 for small study desk and bed store JB 210801 AHFG For families. Provision depends on local operational policies. May be designed for use during the day as overflow interview rooms
| Ensuite - Overnight | OVES | SC | 1 | 10036 | 4 | 4 | 4 |  | 4.2 | 0.2 | 5%
| Ensuite - Overnight | OVES | SC | 1 | 10037 | 4 | 4 | 4 |  | 5 | 1 | 25%
| Bay - Prop | PROP-2 | SC | 1 | 10038 | 2 | 2 | 2 |  | 2.75 | 0.75 | 37.5% |  |  |  | 210801 AHFG Optional. For family / visitor belongings if using overnight stay rooms as multipurpose spaces
| Whanau | PRJ-WH | NS | 1 | 10031 | NA | 18 | 18 |  | 18 |  | 0% |  |  |  | 211030 UG2 room added AG
| Whanau | PRJ-WH | NS | 1 | 10032 | NA | 18 | 18 |  | 22 | 4 | 22.22% |  |  |  | 211030 UG2 room added AG
Sub Total | [5 empty] | 121 |  | 117 |  | 177.35 | 60.35 | 51.58%
Circ. | [5 empty] | 30.25 | 25% | 29.25 | 25% | 44.3375 |  | 51.58%
PATIENT AREAS | Bed - Intensive Care | 1BR-IC | SC | 1 | 20010 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20011 | 25 | 25 | 25 |  | 26 | 1 | 4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20012 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20013 | 25 | 25 | 25 |  | 24 | -1 | -4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20014 | 25 | 25 | 25 |  | 24 | -1 | -4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20015 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20016 | 25 | 25 | 25 |  | 28 | 3 | 12% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20017 | 25 | 25 | 25 |  | 24 | -1 | -4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20018 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20019 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20020 | 25 | 25 | 25 |  | 28 | 3 | 12% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20021 | 25 | 25 | 25 |  | 24 | -1 | -4% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Bed - Intensive Care | 1BR-IC | SC | 1 | 20022 | 25 | 25 | 25 |  | 25 |  | 0% |  | 1 |  | 210801 AHFG The number of enclosed vs open bays provided will depend on the patient cohort and infection control considerations. Includes provision for a palliative care space
| Pt Bay – Intensive Care | PBIC | SC |  | 20023 | 24.5 |  | 0 |  | 0 | [5 empty] | 211029 UG1 number TBC in Concept Design JB
| Bed - Intensive Care (N) | 1BR-IC | SC | 1 | 20024 | 25 | 25 | 25 |  | 24.8 | -0.2 | -0.8% |  | 1 |  | 210801 AHFG No. dependent on patient cohort and local jurisdictional policies
| Bed - Intensive Care (P) | 1BR-IC | SC | 1 | 20025 | 25 | 25 | 25 |  | 24.8 | -0.2 | -0.8% |  | 1 |  | 210801 AHFG Optional. Depends on cohort and local jurisdictional requirements
| Ante | ANRM | SC | 1 | 20026 | 6 | 6 | 6 |  | 6 |  | 0% |  |  |  | 210801 AHFG Attached to 1 Bed Rooms - Isolation Class N (neg pressure ventilation)
| Bath | BATH | SC | 1 | 20027 | 15 | 16 | 16 |  | 15.1 | -0.9 | -5.625% |  |  |  | 210801 AHFG Requires gases. May be shared between 2 pods subject to design
| Ensuite - Special | ENS-SP | SC | 1 | 20028 | 6 | 7 | 7 |  | 7 |  | 0% |  |  |  | 210801 AHFG Ensuites dedicated to negative pressure, positive pressure and standard isolation rooms. Number dependent on isolation room requirements. PICUs should include bathrooms with baths for younger children
| Ensuite - Special | ENS-SP | SC | 1 | 20029 | 6 | 7 | 7 |  | 6.9 | -0.1 | -1.429% |  |  |  | 210801 AHFG Ensuites dedicated to negative pressure, positive pressure and standard isolation rooms. Number dependent on isolation room requirements. PICUs should include bathrooms with baths for younger children
| Ensuite - Special | ENS-SP | SC | 1 | 20030 | 6 | 6 | 6 |  | 6 |  | 0% |  |  |  | 210801 AHFG Ensuites to be shared between remaining rooms (1:4 ratio). PICUs should include bathrooms with baths for younger children
| Ensuite - Special | ENS-SP | SC | 1 | 20031 | 6 | 6 | 6 |  | 5.9 | -0.1 | -1.667% |  |  |  | 210801 AHFG Ensuites to be shared between remaining rooms (1:4 ratio). PICUs should include bathrooms with baths for younger children
| Ensuite - Special | ENS-SP | SC | 1 | 20032 | 6 | 6 | 6 |  | 6.2 | 0.2 | 3.333% |  |  |  | 210801 AHFG Ensuites to be shared between remaining rooms (1:4 ratio). PICUs should include bathrooms with baths for younger children
| Staff Station | SSTN-20 | SC | 1 | 20033 | 25 | 20 | 20 |  | 20.23 | 0.23 | 1.15% |  |  | 5 | 210801 AHFG 1 central staff station per pod
| Office – Clinical | OFF-CLN | SC | 1 | 20034 | 15 | 17 | 17 |  | 14.95 | -2.05 | -12.06% |  |  | 4 | 210801 AHFG 1 per pod
| Office - WS |  | NS | 1 | 20035 | 2 | 1.5 | 1.5 |  | 1.3 | -0.2 | -13.33% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20036 | 2 | 1.5 | 1.5 |  | 1.5 |  | 0% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20037 | 2 | 1.5 | 1.5 |  | 1.8 | 0.3 | 20% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20038 | 2 | 1.5 | 1.5 |  | 1.8 | 0.3 | 20% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20039 | 2 | 1.5 | 1.5 |  | 1.3 | -0.2 | -13.33% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20040 | 2 | 1.5 | 1.5 |  | 1.5 |  | 0% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Office - WS |  | NS | 1 | 20041 | 2 | 1.5 | 1.5 |  | 1.5 |  | 0% |  |  | 1 | 210801 AHFG 1 between 2 enclosed rooms
| Bay - HW - A | BHWS-A | SC | 1 | 20042 | 1 | 1 | 1 |  | 1.2 | 0.2 | 20% |  |  |  | 210801 AHFG Also included within each patient bay
| Bay - HW - A | BHWS-A | SC | 1 | 20043 | 1 | 1 | 1 |  | 0.9 | -0.1 | -10% |  |  |  | 210801 AHFG Also included within each patient bay
| Bay - Linen | BLIN | SC | 1 | 20044 | 2 | 2 | 2 |  | 8.4 | 6.4 | 320% |  |  |  | 210801 AHFG 1 per pod
| Bay - Resus | BRES | SC | 1 | 20045 | 1.5 | 1.5 | 1.5 |  | 6.5 | 5 | 333.3% |  |  |  | 210801 AHFG 1 per pod
| Bay - Path | BPATH | SC | 1 | 20046 | 3 | 3 | 3 |  | 4.9 | 1.9 | 63.33% |  |  |  | 210801 AHFG 1 per pod. May be shared between 2 pods depending on layout
| Bay - PTS | BPTS | SC | 1 | 20047 | 1 | 1 | 1 |  | 1.5 | 0.5 | 50% |  |  |  | 210801 AHFG 1 per pod
| Bay - Equip | BMEQ-4 | SC | 1 | 20048 | 4 | 5 | 5 |  | 6 | 1 | 20% |  |  |  | 210801 AHFG 3 per pod. Locate in quiet low traffic areas
| Bay - Equip | BMEQ-4 | SC | 1 | 20049 | 4 | 5 | 5 |  | 4 | -1 | -20% |  |  |  | 210801 AHFG 3 per pod. Locate in quiet low traffic areas
| Bay - Equip | BMEQ-4 | SC | 1 | 20050 | 4 | 5 | 5 |  | 2 | -3 | -60% |  |  |  | 210801 AHFG 3 per pod. Locate in quiet low traffic areas
| Bay - Blank / Fluid | BBW | SC | 1 | 20051 | 1 | 1 | 1 |  | 1.2 | 0.2 | 20% |  |  |  | 210801 AHFG Shared between 2 pods if collocated, subject to design
Sub Total | [5 empty] | 527 |  | 502 |  | 513.18 | 11.18 | 2.227%
Circ. | [5 empty] | 210.8 | 40% | 200.8 | 40% | 205.272
SUPPORT AREAS | Bay - Bev | BBEV-OP | SC | 1 | 30026 | 4 | 3 | 3 |  | 4.1 | 1.1 | 36.67% |  |  |  | 210801 AHFG Share between 2 pods where collocated
| Bay - Meal Trolley | BMT-4 | SC | 1 | 30027 | 4 | 4 | 4 |  | 3.8 | -0.2 | -5% |  |  |  | 210801 AHFG Optional. Depends on local operational policy for food services. Share between 2 pods where collocated
| Formula |  |  | 0 |  | 9 | 0 | 0 |  | 0 | [5 empty] | 211029 UG21 deleted - not required JB 210801 AHFG Used for formula preparation and storage, and sterilising of expressing equipment within PICUs
| Medication | STDR-10 | SC | 1 | 30029 | 12 | 12 | 12 |  | 15 | 3 | 25% |  |  |  | 210801 AHFG Share between 2 pods where collocated. Area to be distributed if automated dispensing units are provided centrally within central staff area
| Store - Clean |  | NS | 1 | 30030 | 30 | 33 | 33 |  | 27 | -6 | -18.18% |  |  |  | 210801 AHFG Clean stock. Distribution between pods dependent on layout
| Store - Equip | STEQ-20 | SC | 1 | 30031 | 28 | 30 | 30 |  | 35 | 5 | 16.67% |  |  |  | 211029 increased from 112m2 for 3rd access door JB 210801 AHFG 2m2 per bed. May include specialty stores e.g. for retrieval equipment, renal, allied health etc
| Store - General | STGN-9 | SC | 1 | 30032 | 14 | 15 | 15 |  | 14.9 | -0.1 | -0.6667% |  |  |  | 211029 increased from 28m2 for additional shelving UG1 JB
| Respiratory / Biomedical |  | NS | 1 | 30033 | 12 | 35 | 35 |  | 32.3 | -2.7 | -7.714% |  |  |  | 210801 AHFG For set up and servicing
| Dirty Utility | DTUR-12 | SC | 1 | 30034 | 12 | 12 | 12 |  | 12.23 | 0.23 | 1.917% |  |  |  | 210801 AHFG 1 per pod
| Clean-Up | CLUP-7 | SC | 1 | 30035 | 7 | 7 | 7 |  | 7.9 | 0.9 | 12.86%
| Cleaner | CLRM-5 | SC | 1 | 30036 | 5 | 5 | 5 |  | 5.2 | 0.2 | 4% |  |  |  | 210801 AHFG Number and distribution depends on layout
| Disposal | DISP-10 | SC | 1 | 30037 | 8 | 8 | 8 |  | 7.35 | -0.65 | -8.125%
| Simulation |  | NS | 1 | 30038 | 25 | 25 | 25 |  | 23 | -2 | -8% |  |  |  | 220306 room required for educational purposes TK 210801 AHFG Optional. Requirement will depend on education and training requirements / access to central simulation facilities
| Control - Simulation |  |  | 1 |  | 0 | 0 | 0 |  | 0 | [5 empty] | 220306 deleted as not required as less than 56 bed unit TK 210801 AHFG Optional
| Workroom - Telehealth |  | NS | 1 | 30040 | 12 | 12 | 12 |  | 12.1 | 0.1 | 0.8333% |  |  |  | 210801 AHFG Optional. Inclusion depends on operational guideline of unit. Can be used for viewing medical images and discussions, and for telehealth activities for liaison regarding transfers, virtual rounds etc.
| Toilet - Staff | WCST | SC | 1 | 30041 | 3 | 3 | 3 |  | 3.1 | 0.1 | 3.333% |  |  |  | 210801 AHFG For ease of access from clinical areas
| Toilet - Staff | WCST | SC | 1 | 30042 | 3 | 3 | 3 |  | 2.9 | -0.1 | -3.333% |  |  |  | 210801 AHFG For ease of access from clinical areas
| Comms Room | COMM | SC | 1 | 30043 | 0 | 14 | 14 |  | 15 | 1 | 7.143%
| <add additional rooms here> | PRJ-XX
Sub Total | [5 empty] | 188 |  | 207 |  | 220.88 | 13.88 | 6.705%
Circ. | [5 empty] | 75.2 | 40% | 82.8 | 40% | 88.352
STAFF AREAS | Office - 1P | OFF-S9 | SC | 1 | 40001 | 9 | 9 | 9 |  | 10.1 | 1.1 | 12.22% |  |  | 1 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Office - 1P | OFF-S12 | SC | 1 | 40002 | 12 | 12 | 12 |  | 10.95 | -1.05 | -8.75% |  |  | 1 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Office - WS | OFF-WS | NS | 1 | 40003 | 4.4 | 4.4 | 4.4 |  | 4 | -0.4 | -9.091% |  |  | 1 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Office - WS | OFF-WS | NS | 1 | 40004 | 5.5 | 5.5 | 5.5 |  | 5.65 | 0.15 | 2.727% |  |  | 1 | 210801 AHFG Allocation will be dependent on staff profile and jurisdictional policies relating to staff work areas
| Meeting | MEET-L-15 | SC | 1 | 40005 | 15 | 15 | 15 |  | 16.1 | 1.1 | 7.333% |  |  |  | 211029 UG2 increased from 12m2 to include telehealth function JB
| Meeting | MEET-L-20 | SC | 1 | 40006 | 20 | 20 | 20 |  | 18.8 | -1.2 | -6% |  |  |  | 210801 AHFG Assume shared access to 20m2 meeting room under Entry / Waiting & Family Support for smaller services
| Store - Photo /Stationery | STPS-8 | SC-D | 1 | 40007 | 5 | 5 | 5 |  | 14.8 | 9.8 | 196% |  |  |  | 211029 Project wide decision - reduced from 8m2 room to open bay JB
| Staff Room | SRM-18 | SC | 1 | 40008 | 18 | 60 | 60 |  | 50 | -10 | -16.67%
| Change - F | CHST-10 | SC | 1 | 40009 | 10 | 40 | 40 |  | 40 |  | 0% |  |  |  | 210801 AHFG Includes toilets, showers, lockers; size depends on the staffing per shift. (Female)
| Change - M | CHST-10 | SC | 1 | 40010 | 8 | 30 | 30 |  | 32 | 2 | 6.667% |  |  |  | 211029 reduced from 32m2 less lockers required JB 210801 AHFG Includes toilets, showers, lockers; size depends on staffing per shift. (Male)
| Toilet - Access | WCAC | SC | 1 | 40011 | 6 | 6 | 6 |  | 6.3 | 0.3 | 5%
| Bed - Overnight | OVBR | SC | 1 | 40012 | 10 | 10 | 10 |  | 12 | 2 | 20% |  |  |  | 210801 AHFG Optional. Requirement depends on staffing arrangements. Close access to toilet / shower facilities required
Sub Total | [5 empty] | 122.9 |  | 216.9 |  | 220.7 | 3.8 | 1.752%
Circ. | [5 empty] | 30.725 | 25% | 54.225 | 25% | 55.175
CORRIDORS | Corridor | PRJ-COR | SC-D | 1 | 60002 |  |  |  |  | 42 | [5 empty] | *add designed area only - leave brief m2 blank as this is breifed as a %*
| Corridor | PRJ-COR | SC-D | 1 | 60005 |  |  |  |  | 60 | [5 empty] | *add designed area only - leave brief m2 blank as this is breifed as a %*
| Corridor | PRJ-COR | SC-D | 1 | 60008 |  |  |  |  | 35 | [5 empty] | *add designed area only - leave brief m2 blank as this is breifed as a %*
PLANT | DB board | PRJ-PLT | NS | 1 | 70001 |  |  |  |  | 2 | [5 empty] | *i.e comms, FHR, risers, small localised plant room - leave brief m2 blank as this is breifed as a % (excluding Comms Room)*
| Service Riser 1 | PRJ-PLT | NS | 1 | 70002 |  |  |  |  | 2.5 | [5 empty] | *i.e comms, FHR, risers, small localised plant room - leave brief m2 blank as this is breifed as a % (excluding Comms Room)*
| FHR (fire hose reel) | PRJ-PLT | NS | 1 | 70003 |  |  |  |  | 0.5
| Plant Room | PRJ-PLT | NS | 1 | 70004 |  |  |  |  | 150
Total Circ |  |  |  |  | AHFG | 346.975 | Brief | 367.075 | Designed | 137 | 128.42 | -62.68%
Total Area - Rooms |  |  |  |  | AHFG | 958.9 | Brief | 1042.9 | Designed | 1132.11 | 256.84 | 8.554%
Grand Total | [5 empty] | 1305.875 |  | 1409.975 |  | 1269.11 | 226.21 | -9.991%
Total Plant 11% | [6 empty] | Brief | 155.0973 | Designed | 155 | -0.09725 | -0.0627% |  |  |  | 11% of Grand Total - GFA
Functional Unit Total | [12 empty] | 0 | 15 | 20
EXTERNAL | Terrace - ICU Bed |  | NS | 1 | 90000 | 0 | 50 | 50 |  | 45 | -5 | -10% |  |  |  | 210801 Added. Ability for patients to have fresh air and light JC
| Staff Outdoor area |  | NS | 1 | 90001 | 0 | 0 | 0 |  | 20 | 20 |  |  |  |  | 210912 Added. Staff relaxation area JC
Grand Total | [7 empty] | 50 |  | 65 | 65 | 30%

## AreaSUM Example (hidden)
OVERALL AREA COMPARISON
Briefed Vs. Designed Building/Department Summaries
Issue Date: |  | 2021-10-28
HEALTH PLANNING UNITS |  | BRIEFED |  | DESIGNED vs BRIEFED INC. CIRC |  |  | CONCEPT vs BRIEFED |  |  | PRELIM vs BRIEFED INC. CIRC |  |  | DEVELOPED vs BRIEFED INC. CIRC |  |  | DETAILED vs BRIEFED INC. CIRC |  |  | CONSTRUCTN vs BRIEFED INC CIRC |  |  | NOTES | FUNCTIONAL UNIT COUNT
|  | Briefed Area | Briefed Inc Circ. | Designed Area | Diff. m2 | Diff. % (+/- 5%) | Concept Area | Diff. m2 | Diff. % (+/- 5%) | Prelim Area | Diff. m2 | Diff. % (+/- 5%) | Developed Area | Diff. m2 | Diff. % (+/- 5%) | Detailed Area | Diff. m2 | Diff. % (+/- 5%) | Construction Area | Diff. m2 | Diff. % (+/- 5%) |  | Patient Spaces | Procedure Rooms | Staff W/Station
Intensive Care Unit |  | 1042.9 | 1409.975 | 1269.11 | -140.865 | -9.991% | 1499 | 89.025 | 6.314% | 1495 | 85.025 | 6.03% | 1497 | 87.025 | 6.172% | 1497 | 87.025 | 6.172% | 1497 | 87.025 | 6.172% | e.g. incl 250m2 unallocated | 12 |  | 10
Emergency Department |  | 3000 | 3000 | 3370 | 370 | 12.33% | 3360 | 360 | 12% | 3330 | 330 | 11% | 3338 | 338 | 11.27% | 3335 | 335 | 11.17% | 3336 | 336 | 11.2% |  | 50 | 4 | 20
Radiology |  | 2000 | 2000 | 1899 | -101 | -5.05% | 1897 | -103 | -5.15% | 1905 | -95 | -4.75% | 1930 | -70 | -3.5% | 1932 | -68 | -3.4% | 1935 | -65 | -3.25% |  |  | 5 | 12
Surgical Ward 1 |  | 1500 | 1500 | 1507 | 7 | 0.4667% | 1510 | 10 | 0.6667% | 1535 | 35 | 2.333% | 1535 | 35 | 2.333% | 1536 | 36 | 2.4% | 1536 | 36 | 2.4% |  | 32 |  | 10
Surgical Ward 2 |  | 1500 | 1500 | 1506 | 6 | 0.4% | 1501 | 1 | 0.06667% | 1515 | 15 | 1% | 1515 | 15 | 1% | 1515 | 15 | 1% | 1514 | 14 | 0.9333% |  | 32 |  | 10
Front of House |  | 400 | 435 | 300 | -135 | -31.03% | 350 | -85 | -19.54% | 360 | -75 | -17.24% | 375 | -60 | -13.79% | 372 | -63 | -14.48% | 376 | -59 | -13.56%
Theatres |  | 2000 | 2050 | 2000 | -50 | -2.439% | 1999 | -51 | -2.488% | 2005 | -45 | -2.195% | 2005 | -45 | -2.195% | 2005 | -45 | -2.195% | 2004.5 | -45.5 | -2.22% |  |  | 5
PACU |  | 200 | 250 | 260 | 10 | 4% | 300 | 50 | 20% | 301 | 51 | 20.4% | 301.5 | 51.5 | 20.6% | 302 | 52 | 20.8% | 302 | 52 | 20.8% |  | 15 |  | 5
Admissions / Stepdown |  | 300 | 330 | 325 | -5 | -1.515% | 315 | -15 | -4.545% | 312 | -18 | -5.455% | 311 | -19 | -5.758% | 311 | -19 | -5.758% | 311 | -19 | -5.758% |  | 20 |  | 10
Helipad |  | 200 | 250 | 200 | -50 | -20% | 205 | -45 | -18% | 202 | -48 | -19.2% | 206 | -44 | -17.6% | 219 | -31 | -12.4% | 219 | -31 | -12.4%
Travel | 18% | 2185.722 | 2290.4955 | 2274.4998 | -15.9957 | -0.6984% | 2328.48 | 37.9845 | 1.658% | 2332.8 | 42.3045 | 1.847% | 2342.43 | 51.9345 | 2.267% | 2344.32 | 53.8245 | 2.35% | 2345.49 | 54.9945 | 2.401%
Enginering | 25% | 3035.725 | 3181.2438 | 3159.0275 | -22.2162 | -0.6984% | 3234 | 52.7562 | 1.658% | 3240 | 58.7562 | 1.847% | 3253.375 | 72.1312 | 2.267% | 3256 | 74.7562 | 2.35% | 3257.625 | 76.3812 | 2.401%
Façade | 0% | 0 | 0 | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0 |  |  | 0
Net |  | 12142.9 | 12724.975 | 12636.11 | -88.865 | -0.6984% | 12936 | 211.025 | 1.658% | 12960 | 235.025 | 1.847% | 13013.5 | 288.525 | 2.267% | 13024 | 299.025 | 2.35% | 13030.5 | 305.525 | 2.401% |  | Total Pt Spaces | Total Proc Rms | Total Staff W/Stations
Gross |  | 17364.347 | 18196.7143 | 18069.6373 | -127.0769 | -0.6984% | 18498.48 | 301.7657 | 1.658% | 18532.8 | 336.0857 | 1.847% | 18609.305 | 412.5907 | 2.267% | 18624.32 | 427.6057 | 2.35% | 18633.615 | 436.9007 | 2.401% |  | 161 | 14 | 77
COMPARISONS: | [21 empty] | NOTES
HPU 7-9 e.g. Perioperative Floor GDA |  | 2500 | 2630 | 2585 | -45 | -1.711% | 2614 | -16 | -0.6084% | 2618 | -12 | -0.4563% | 2617.5 | -12.5 | -0.4753% | 2618 | -12 | -0.4563% | 2617.5 | -12.5 | -0.4753% | Includes Theatres, PACU, Admissions/ Stepdown
HPU 4-5 e.g. 2x Surgical Wards GDA |  | 3000 | 3000 | 3013 | 13 | 0.4333% | 3011 | 11 | 0.3667% | 3050 | 50 | 1.667% | 3050 | 50 | 1.667% | 3051 | 51 | 1.7% | 3050 | 50 | 1.667% | Includes 2x surgical wards